##### May 2008

This distribution contains change pages for patch MD\*1.0\*6 of the Clinical Procedures 1.0 Implementation Guide.

The change pages for CP Patch 4 and CP Patch 14 should be inserted before the change pages for CP Patch 6:

File Name: Patch:

MD\_1\_P4\_IMPG.PDF MD\*1.0\*4

MD\_1\_P14\_IMPG.PDF MD\*1.0\*14

Patch MD\*1.0\*6 pages:

Replace Pages: With Pages:

Title page Title page

Revision History Revision History

Table of Contents Table of Contents

1-9 to 1-12 1-9 to 1-12

6-3 to 6-6 6-3 to 6-6

6-11 to 6-30 6-11 to 6-30

11-1 to 11-12 11-1 to 11-12

13-3 to 13-4 13-3 to 13-4

14-3 to 14-6 14-3 to 14-6

16-1 to 16-6 16-1 to 16-6

17-1 to 17-2 17-1 to 17-2

18-1 to 18-12

Index Index

**Note:** A new chapter “Appendix D – Exported Values For Hemodialysis Options” was added before the Index, which was Chapter 18. The Index is now Chapter 19.



**CLINICAL PROCEDURES IMPLEMENTATION GUIDE**

Version 1.0

April 2004

Revised May 2008

Department of Veterans Affairs Health Systems Design and Development

Provider Systems

# Revision History

|  |  |  |
| --- | --- | --- |
| **Description** | **Date** | **Technical Writer** |
| Originally released. | April 2004 |  |
| [1](#_bookmark0)Patch MD\*1.0\*4 released. | September 2006 | REDACTED |
| [2](#_bookmark1)Patch MD\*1.0\*9 released November 2007. Update Setting up HL7 Parameter for port 5000 with CACHE. | February 2008 | REDACTED |
| [3](#_bookmark2) Patch MD\*1.0\*14 released. Updated Setting Up Consults for Clinical Procedures, Exported XPAR Kernel Parameters, add new section called Scheduled Options. Added information about launching CP Gateway under the section Working with CP Gateway. | March 2008 | REDACTED |
| [4](#_bookmark3)Patch MD\*1.0\*6 released. Updated  MD namespace Clinical Procedures file list and CP Class Upload Header output display, added TIU prompts for adding new TIU Note Titles, added instrument warning for automated instruments, added Processing Application field, changed wording for Count/Non-count clinics, added new Exported Kernel XPAR parameters and screen capture, revised “Setting Up HL7 Parameters chapter for clarity, updated list of Instrument Processing Routines, added Appendix D – Exported Values For Hemodialysis Options. | May 2008 | REDACTED |

1 Patch MD\*1.0\*4 September 2006 Patch 4 release added.

2 Patch MD\*1.0\*9 November 2007 Patch 9 release added.

3 Patch MD\*1.0\*14 March 2008 Patch 14 release added.

4 Patch MD\*1.0\*6 May 2008 Patch 6 release added.

# Table of Contents

1. Introduction 1-1

About Clinical Procedures 1-1

About CP User 1-8

About CP Manager 1-8

About CP Gateway 1-8

Intended Audience 1-8

[Related Manuals 1-9](#_bookmark4)

[General CP Package Information 1-10](#_bookmark6)

[Resource Requirements 1-11](#_bookmark8)

[Hospital Location File Requirement (Implementing Workload Reporting) 1-12](#_bookmark9)

[VistA Imaging 1-12](#_bookmark9)

1. Using CP Manager 2-1

CP Manager Toolbar 2-1

Finding a Parameter 2-2

Deleting an Automated Instrument or Procedure 2-3

Printing Reports 2-4

1. About Test Accounts and Imaging 3-1

Changing All Test Accounts 3-2

Changing the Current Namespace 3-2

Configuring the Imaging Display Station 3-5

Changing Test Accounts that Use a Background Processor 3-6

Connecting the Background Processor PC to VistA Servers 3-9

Refreshing Existing Test Accounts 3-10

1. Setting Up TIU for Clinical Procedures 4-1

Step 1 - Verify Clinical Procedures Class Upload Header 4-2

Step 2 - Create CP Class Document Definitions 4-3

Example of New TIU Prompts 4-9

Step 3 - Define Clinical Procedures Class Document Parameters 4-10

1. About ASU Business Rules and the Role of the Interpreter 5-1

How Business Rules Work 5-1

Role of the Interpreter 5-3

1. Setting Up Clinical Procedures 6-1

Step 1 - Populate the CP Definition (#702.01) file 6-1

Step 2 – Setting Up Instruments 6-2

[Editing an Automated Instrument 6-3](#_bookmark10)

Adding an Automated Instrument 6-8

[Using the Instrument Analyzer 6-11](#_bookmark12)

[Step 3 – Setting Up Procedures 6-12](#_bookmark14)

[Editing a Procedure 6-12](#_bookmark14)

[Adding a Procedure 6-16](#_bookmark21)

[Step 4 – Setting Up System Parameters 6-18](#_bookmark25)

[Allow non-instrument attachments 6-19](#_bookmark27)

[Bypass CRC Checking 6-20](#_bookmark30)

[Clinical Procedures Home Page 6-20](#_bookmark30)

[Clinical Procedures On-Line 6-20](#_bookmark30)

[CP/BGP Transfer Directory 6-20](#_bookmark30)

[CRC Values 6-21](#_bookmark31)

[Calculating a File’s CRC Value 6-22](#_bookmark34)

[Days to keep instrument data 6-22](#_bookmark34)

[Imaging File Types 6-23](#_bookmark35)

[Offline Message 6-24](#_bookmark38)

[Version Compatibility 6-24](#_bookmark38)

[VistA Scratch HFS Directory 6-26](#_bookmark43)

[Step 5 – Exported Kernel XPAR Parameters 6-27](#_bookmark44)

[Exported Kernel XPAR Parameters for Patch MD\*1.0\*14 6-27](#_bookmark44)

[Exported Kernel XPAR Parameters for Patch MD\*1.0\*6 6-29](#_bookmark46)

1. Scheduled Options 7-1
2. Setting Up Consults for Clinical Procedures 8-1

Step 1 – Setting Up Consult Services 8-1

Step 2 - Creating Consult Procedures 8-5

1. Setting Up CPRS for Clinical Procedures 9-1

Step 1 – Setting Up the Notification 9-1

Step 2 – Editing Parameters 9-3

Ask Encounter Update (ORWPCE ASK ENCOUNTER UPDATE) 9-4

Broadcast Messages to Other Apps (ORWOR BROADCAST MESSAGES) 9-5

Force PCE Entry (ORWPCE FORCE PCE ENTRY) 9-6

Add CP User to the CPRS Tools Menu (ORWT TOOLS MENU) 9-7

1. Working with CP Gateway 10-1

Log File Options 10-4

1. [Setting Up HL7 Parameters 11-1](#_bookmark48)

[Configuration Instructions Information 11-1](#_bookmark48)

[IP Addresses and Ports 11-2](#_bookmark51)

[Setting Up a New HL7 Single Listener for High-Volume Devices 11-2](#_bookmark51)

[Creating a Logical Link 11-3](#_bookmark54)

[Creating a Device Protocol Client 11-5](#_bookmark56)

[Activating the Logical Links 11-7](#_bookmark58)

[Adding a Device Client as a Server Subscriber 11-7](#_bookmark58)

[Using Port 5000 11-9](#_bookmark61)

[Benefits of Using a Single Port Listener 11-9](#_bookmark61)

[Setting Up Port 5000 11-9](#_bookmark61)

[File Settings 11-9](#_bookmark61)

[Technical Issues 11-11](#_bookmark67)

1. Configuring the Automated Instrument Share Folder 12-1
2. Troubleshooting 13-1

ii Clinical Procedures V. 1.0 April 2004

1. Glossary 14-1
2. Appendix A – CP Application Startup Options and Command Line Switches 15-1

Introduction 15-1

What is a Command Line Switch? 15-1

Shared Broker Environment 15-1

CPRS Tools Menu 15-2

All Command Line Switches 15-3

1. [Appendix B – Exported Procedures List 16-1](#_bookmark70)
2. [Appendix C - Instrument Processing Routines 17-1](#_bookmark71)
3. [Appendix D – Exported Values For Hemodialysis Options 18-1](#_bookmark74)

[Custom Data List 18-1](#_bookmark74)

[Anticoagulants 18-1](#_bookmark74)

[Code Statuses 18-1](#_bookmark74)

[Dialyzer List 18-1](#_bookmark74)

[Education Codes 18-3](#_bookmark76)

[ESRD Diagnosis 18-3](#_bookmark76)

[Medication Routes 18-5](#_bookmark77)

[Medication Units 18-5](#_bookmark77)

[Modalities 18-6](#_bookmark78)

[TIU Note Titles 18-6](#_bookmark78)

[Transportation Methods 18-6](#_bookmark78)

[Preferences 18-7](#_bookmark79)

[System Preferences 18-7](#_bookmark79)

[Report List 18-8](#_bookmark80)

[Summary Report Template 18-8](#_bookmark80)

1. [Index 19-1](#_bookmark81)

April 2004 Clinical Procedures V. 1.0 iii

## Related Manuals

Here is a list of related manuals that you may find helpful: Clinical Procedures Installation Guide

Clinical Procedures Technical Manual and Package Security Guide

Clinical Procedures User Manual Clinical Procedures Release Notes CPRS User Guide: GUI Version CPRS Setup Guide

Consult/Request Tracking User Manual Consult/Request Tracking Technical Manual

Text Integration Utilities (TIU) Implementation Guide Text Integration Utilities (TIU) User Manual

VistA Imaging System (Clinical) User Manual

These manuals can be found in the [VistA Documentation Library (VDL)](http://www.va.gov/vdl/), <http://www.va.gov./vdl>. Select **Clinical** from the VDL web page, select the package you want, and then select the manuals. For example, you can select CPRS on the left side of the page. The list of CPRS manuals is displayed.

[1](#_bookmark5) You may also want to read the CP Implementation Process (Webpage), which is available on the CP website. Go to <http://vista.med.va.gov/ClinicalSpecialties/clinproc/>.

Point to **Clinical Procedures Project,** then click **Documentation.** When the Documentation page displays, click **Clinical Procedures Documents,** then click **The CP Implementation Process (Webpage).** This list includes a high-level step-by-step guide to the installation and the implementation process.

1 MD\*1.0\*6 May 2008 Changed document name from “Site Installation Checklist” to “The CP Implementation Process (Webpage).” Revised directions to access the document.

## General CP Package Information

1. Name spacing and file listing.

Clinical Procedures is found in the MD namespace. All routines, templates and options begin with MD. File numbers range from 702 to 704 and are stored in the ^MDD and ^MDS globals. The range of 704.201 to 704.209 is stored in the ^MDK global. Here is a list of the Clinical Procedures files:

#702 CP Transaction

#702.01 CP Definition

#702.09 CP Instrument #703.1 CP Result Report

[1](#_bookmark7)#703.9 CP Conversion File

#704.201 Hemodialysis Access Points File #704.202 Hemodialysis Study File #704.209 Hemodialysis Setting File

1. Queuing TaskMan jobs.

Queued TaskMan jobs are not associated with this application.

1. Accessing modules.
   * Assign the option [MD GUI USER] to the clinical staff, who need access to CP User.
   * Assign the option [MD GUI MANAGER] to the Clinical Application Coordinator, CP package coordinator, and Information Resource Management Service (IRMS) staff for access to CP Manager.
   * Assign the MD MANAGER key to the Clinical Application Coordinator or the CP Package Coordinator. This key controls access to the Update Study Status menu option that allows clinicians to fix study errors. This key also controls access to the Delete Study option.
   * Assign the MAGCAP CP user security key to technicians, who will be using VistA Imaging to capture a consent form and link it to a CP study or TIU document.
2. Printer issues.

All reports are printed to Client (Windows) printers.

1. Online Help.

1 Patch MD\*1.0\*6 May 2008 Files added.

Online help is available when questions arise. Click Help or choose Help from the menu bar. You can also press F1 for help on a specific window.

1. Automatic Version Updates.

CP applications (client and server) do not contain automatic update capabilities. You must remove the previous version before you can install the new version.

1. Command line switches.

For alternate methods of running Clinical Procedures, refer to Appendix A - CP Application Startup Options and Command Line Switches, p. 9-7.

## Resource Requirements

* + Clinical Procedures can only run at sites that are running VistA Imaging V. 3.0.
  + Workstations must run Windows 2000 or later. 12 MB of available disc space is required.

VistA Server resources:

Globals Type of Data Size

^MDS Static global 25 k

^MDD Patient data for the 25-75 k/patient Clinical Procedures

^MDK Hemodialysis Studies 25-75 k/patient

##### NOTE: These globals must all be journaled.

## Hospital Location File Requirement (Implementing Workload Reporting)

Be sure that the hospital location entry (Hospital Location #44 file) for each CP procedure contains the correct Institution field entry. The Hospital Location is used for workload reporting. (The Institution field tells VistA Imaging where to store the images on the server. If there is no Institution field, CP defaults to the institution of the user who logged on to CP Gateway.)

## VistA Imaging

Providers at a site must use the VistA Imaging Display client to view CP results and reports. Be sure that VistA Imaging V.3.0 or greater and Patch 7 of Imaging V.3.0 (MAG\*3.0\*7) are installed.

In most cases, you edit an existing automated instrument. The Mallinckrodt Clinivision, Olympus Endoworks, GE Medical Systems Muse and Viasys/Sensormedics Vmax automated device interfaces are exported with Clinical Procedures. You must edit all the automated instruments that you want to implement with necessary information.

### Editing an Automated Instrument

The following list of fields applies to automated instruments:

\* indicates fields that must be filled in for an active instrument to work properly.

**BOLD** indicates fields that are already populated when an automated instrument is exported.

##### Instrument Name

* **Printable Name Description**
* **M Routine**
* **Pkg. Code**
* **Valid Attachment Types**

**If Bi-Directional Instrument is checked:**

**HL7 Inst ID**

**HL7 Link**

* Notification Mailgroup
* Active

Serial Number (Optional)

Delete When Submitted (Optional) Default Extension (Optional)


IP Address (Optional) Port (Optional)

HL7 Unv Svc ID (Optional) Server Name (Optional) Share Name (Optional) Path Name (Optional)

Executable Name (Optional)

To edit an automated instrument:

1. View the list of automated instruments. See [Figure 6-2](#_bookmark11).
2. Click on an automated instrument. The edit screen is displayed on the right side of the Clinical Procedures Manager window.
3. Enter the fields that apply to the instrument you selected.
4. Click **Save** when you are done.
5. Click **Print** if you want to print an Automated Instrument report. See Printing Reports, p. 2-4.



**Figure 6-2**

Here is a list of fields for automated instruments.

**General**: This section contains general information about the instrument.

**Instrument Name:** If you are editing an instrument, the name is filled in.

**Note:** This field must be filled in for an active instrument to work properly.

If you are adding a new instrument that is already supported by CP, do one of the following:

* + If the device is bi-directional, you can enter a name of your own choice (3-30 characters), such as Muse EKG (Tampa). The name does not have to be the vendor’s name.
  + If the device is uni-directional, you need to enter a CP defined name. In this case, you can contact TSO or NVS for the correct instrument name.

If you are adding a new instrument (bi-directional or uni-directional) that is not supported by CP, then you must enter a NOIS/Remedy help ticket. Keep in mind that adding unsupported instruments is a complex task and may cause some image quality problems.

**Printable Name**: Enter a name for the instrument report (3 to 30 characters). You can use the same name as the instrument name. This name is used as the printable name on reports. Must be filled in for an active instrument to work properly.

**Notification Mailgroup**: Enter the name of a local VistA mailgroup that contains a list of people, who will be notified if a problem arises with this automated instrument.

CP also exports a mailgroup called “MD DEVICE ERRORS” that can be used to populate this field. Enter MD and the field fills in with “MD DEVICE ERRORS’. The coordinator of this group is assigned during package installation. Must be filled in for an active instrument to work properly.

**Description**: Enter a description of the automated instrument (1-50 characters). Optional.

**Serial Number**: Enter the serial number of the automated instrument (1-50 characters). The serial number is used for documentation purposes. Optional.

**Active**: Select this option if you want to make the instrument active and able to transmit results. Do not select if the package coordinator wants to prevent data from a specific automated instrument from being processed. A package coordinator may want to enter the basic information for an automated instrument and not make it active. Must be selected to make this instrument active.

**Attachment Processing**: This section contains information about attachments.

**Delete when Submitted**: Select this option if the medical center does not want to store a duplicate report outside of Imaging, or if the vendor wants to delete files because of storage issues. The vendor determines whether or not the report files can be deleted.

This information is found in the vendor’s setup instructions. Optional.

**Default Ext.**: Enter a default file extension that is exported by the vendor, such as .html,

.jpg, and .pdf. This information should be obtained from the vendor or will be exported with future patches. Optional.

**M Routine**: Indicates the M routine used to process the HL7 message from the automated instrument (1-8 characters). Enter an M routine if the site is entering a new device. The routine must have a namespace of MDHL7\* for any nationally released interfaces. This field also is automatically populated when an instrument interface patch is installed. If a local M routine is developed, use the local namespace. Refer to [Appendix C – Instrument Processing Routines](#_bookmark71), p. [17-1](#_bookmark72), for a list of appropriate M routines for each instrument. Must be filled in for an active instrument to work properly.

**Pkg. Code**: Indicates which package is to process the instrument results. Must be filled in for an active instrument to work properly.

Medicine: Select if your study data is stored in the Medicine package. If a site is currently running Medicine and has an instrument used for Medicine, you can send the result to Medicine by selecting this field.

CP V1.0 Select if your study data is stored as a final report (in the format of an Imaging document) in Clinical Procedures.

**Valid Attachment Types**: Data types let CP know what kind of data output to expect from the automated instrument so that the data can be processed by the interface routines. The vendor setup instructions provide this information, or Clinical Procedures automatically exports this information. Must be filled in for an active instrument to work properly.

Here is a list of valid attachment types:

UNC (**U**niversal **N**aming **C**onvention or **U**niform **N**aming **C**onvention) - A PC format for specifying the location of resources on a local-area network (LAN).

UUENCODE (**U**nix-to-**U**nix **ENCODE**) - A set of algorithms for converting files into a set of ASCII characters that can be transmitted over a network.

Text - Text stored as ASCII codes.

XML (e**X**tensible **M**arkup **L**anguage) - A specification developed by the World Wide Web Consortium (W3C), the organization that sets standards for the web. XML is a pared-down version of SGML. Designed especially for Web documents.

URL (**U**niform **R**esource **L**ocator) - The global address of documents and other resources on the World Wide Web.


XMS - An XML Style Sheet.

DLL (**D**ynamic **L**ink **L**ibrary) - A library of executable functions or data that can be used by a Windows application.

### Using the Instrument Analyzer

Use the Instrument Analyzer to see if an automated instrument is ready to use with CP.

1. Select **Tools > Instrument Analyzer**.
2. Select the instrument that you want to analyze. Click **Analyze**. A window similar to [Figure 6-6](#_bookmark13) is displayed. This window indicates the ready status of the instrument and lists other information as well.



**Figure 6-6**

* + Ready Status - Pass or Fail. If the Ready Status is Fail, a list of missing fields for that automated instrument is displayed.
  + If an Imaging share directory has not been configured, the following message is displayed “No Imaging Share indicated in the System Parameters.”
  + If the M Routine (processing routine) is not in the MD or MC namespace, a warning is displayed indicating that the M Routine is not in the package namespace.

1. Click **Print** or **Close**.

## Step 3 – Setting Up Procedures

Information on procedures is **not complete** after populating the CP Definition file. **You must go into CP Manager and enter the necessary fields before the package will work successfully.**

If the INIT^MDPOST routine was run, a limited number of exported procedures are stored in a subfolder called Unassigned within the Procedures folder. If the INIT^MDPOST routine was not run, then you need to add new procedures. Since all procedures are initially inactive, you need to activate existing procedures and associate them with treating specialties.

### Editing a Procedure

If the procedures have been exported, then you can edit them as needed. Using CP Manager, you must move each procedure that you want to activate from the Unassigned folder to a treating specialty folder.

* + Double-click the procedure. Now you can edit the procedure, complete the necessary fields, and make the procedure active.
  + To activate the procedure, be sure to select the Active field, and then fill in the following fields to ensure that the procedure works properly

Treating Specialty TIU Note Title Hospital Location

To edit a procedure:

1. [View the list of procedures. See Figure 6-7](#_bookmark17).
2. Click a procedure name. The edit screen is displayed on the right side of the Clinical Procedures Manager window.
3. Enter the fields as applicable.


[**1**](#_bookmark16)**Note:** Make sure to set the Processing Application field to **HEMODIALYSIS** for Hemodialysis procedures.

1. Click **Save** when you are done.

1 Patch MD\*1.0\*6 May 2008 Processing Application field added.

1. If you selected a different treating specialty folder, a confirmation message is displayed. Click **OK** to confirm that the procedure is in the correct treating specialty folder.
2. Click **Print** if you want to print a Procedure report. See Printing Reports, p. 2-4.

**Note:** A procedure can only be deleted through the main menu bar. Refer to the section Deleting an Automated Instrument or Procedure, p. 2-3, for more information. If a procedure has been assigned through Consults, it cannot be deleted.



**[1](#_bookmark18)Figure 6-7**

Here is a list of fields for Procedures.

**General**: This section contains general information about the procedure.

**Procedure Name**: Enter a name used to uniquely identify the procedure (3-30 characters). It is recommended that you enter the name in uppercase, such as PACEMAKER FOLLOWUP.

After you complete the edits, if you entered the name in upper case, the procedure name that you just entered is displayed in title case, Pacemaker Follow-up, (the first letter of every word [is capitalized), in the left side of the CP Manager window. See Figure 6-6](#_bookmark13).

**Active**: Select if you want the procedure to be mapped to Consults. Only active procedures can be selected and linked to the Consults package. Be sure to fill in the Treating Specialty, TIU Note Title, and Hospital Locations fields. Do not select if you do not want procedures to display. Must be selected to make this procedure active.

**Treating Specialty**: Enter at least two letters of a treating specialty, such as CA for CARDIOLOGY, and then click the down arrow to select an appropriate match from the list. This list comes from the Treating Specialty (#45.7) file. Must be filled in for an active procedure to work properly.

**TIU Note Title**: Enter at least two letters of a TIU Note Title, such as CP CARD for CP CARDIOLOGY NOTE or CARD for CARDIOLOGY, and then click the down arrow to select an appropriate match from the list, which comes from the 8925.1 file. This title must be in the CLINICAL PROCEDURES CLASS. Must be filled in for an active procedure to work properly.

[**1**](#_bookmark19)**Hospital Location**: Enter at least two letters of a hospital location, such as CA for Cardiac Clinic, and then click the down arrow to select an appropriate match from the list, which comes from the #44 file. The Hospital Location file is the location where the workload credit for the procedure is tracked and is needed so CPRS can display the appropriate encounter form when prompted. Must be filled in for an active procedure to work properly.

You can enter a COUNT or NON-COUNT clinic for the hospital location.

* + A COUNT clinic captures workload. Patients must be checked in and checked out and an encounter form must be completed in order to collect workload.
  + A NON-COUNT clinic is used only for scheduling purposes and not for workload reporting.

There are three options available for setting up your clinics. The appropriate option for your site depends on how you currently do business and should be discussed with your project implementation manager.

* + COUNT clinic for scheduling purposes / NON-COUNT clinic for CP User. Patient must be checked in/out and encounter form completed on the scheduled appointment. CP User appointment will not collect workload.
  + NON-COUNT clinic for scheduling purposes / COUNT clinic for CP User. Appointment in scheduling package does not need to be checked in/out, nor does an encounter form need to be completed for the appointment. The check in/out and encounter form must be completed for the appointment created through CP User.

1 Patch MD\*1.0\*4 September 2006 Wording for Count/Non-count clinic modified.

* + COUNT clinic for scheduling purposes that passes over to CP User. Patient must be checked in/out and encounter form must be completed. Note, however, that if you use Appointment Manager to check in the patient, you may have to wait up to thirty minutes before you can check-in the patient to CP. During the thirty-minute timeframe, the Patient Care Encounter (PCE) application establishes the visit date. (If you use the Scheduling application to capture workload, make sure that the clinic location is the same as the default location in the Hospital Location field.)

**Auto Submit to VistA Imaging**: Select if a procedure is processed by a bi-directional instrument and additional data does not need to be matched. The study is automatically submitted to V*IST*A Imaging. If this field is not selected, the study will be in the Ready to Complete status. Optional.

**Require External Data**: Select if you want this procedure to allow external attachments. For example, you might want to attach an independent report from a VA or non-VA health care facility. If you want to manually select external attachments, you must select this field.

Be sure the **Allow Non-Instrument Attachments** checkbox is selected in **CP Manager > System Parameters**. There is no default for this field.

**External Attachment Directory**: If you select Require External Data, enter the path where the data is located, or browse to locate a directory (3-150 characters). There is no default on this field. You can locate any directory on the LAN. This is the directory that CP User accesses to find attachments. This directory must be a network share directory that the VistA Imaging Background Processor can access.


[**1**](#_bookmark20)**Processing Application:** Set the Processing Application field to HEMODIALYSIS for Hemodialysis procedures. Any other CP procedures will default to the Default setting, so you do NOT need to set the field.

**Allowable Instruments**: Select each automated instrument that provides results for this procedure. You can select more than one instrument for a procedure. If you only want to use external attachments, do not select any instruments.

You can select both **Allowable** Instruments and **Require External Data**. For example, you can have a pathology report from an endoscopy and you can attach the report to the procedure.

### Adding a Procedure

Before you add a procedure, you can check to see if an appropriated titled procedure already exists that meets your needs. To view the names of procedures, select Procedures and then the appropriate treating specialty folder. A list of procedures is displayed. See [Figure 6-8](#_bookmark22).




**Figure 6-8**

Active procedure icon - Identifies an active procedure Inactive procedure icon - Identifies an inactive procedure

If you decide that you do need to add a procedure, follow these instructions:

1. Select **File > New > Procedure**.
2. Enter the name of the procedure that you want to add. It is recommended that you enter the name in uppercase with a minimum of 3 characters and a maximum of 30 characters.
3. Click **OK**. The Edit screen is displayed. [Figure 6-9](#_bookmark23) is the edit screen for procedures. The Procedure Name that you just entered is displayed in the left side of the CP Manager window in the Unassigned folder.
4. Enter data for each field as applicable. Refer to [Editing a Procedure](#_bookmark14), p. [6-12](#_bookmark15), for detailed field descriptions.
5. Click **Save** when you are done. After you complete the edits, if you entered the name in upper case, the procedure name that you just entered is displayed in title case.
6. Click **OK.** The new procedure appears in the list on the left side of the CP Manager window. Check that the procedure is placed in the correct treating specialty folder.
7. Click **Print** if you want to print a Procedure report. See Printing Reports, p. 2-4.



**[1](#_bookmark24)Figure 6-9**

## Step 4 – Setting Up System Parameters

System parameters are system-wide and affect all procedures and instruments. You must select Clinical Procedure On-Line, and fill in the Imaging Network Share and the VistA Scratch HFS Directory fields for CP to work properly. You can edit the other parameters as required for your site.

Here is a list of the system parameters:

\* Indicates fields that must be filled in for CP to work properly. [Allow non-instrument attachments](#_bookmark27)

[Bypass CRC Checking](#_bookmark30)

[Clinical Procedures Home Page](#_bookmark30)

\*[Clinical Procedures On-Line](#_bookmark30) [1](#_bookmark26)\* [CP/BGP Transfer Directory](#_bookmark30) [CRC Values](#_bookmark31)


[Days to keep instrument data](#_bookmark34) [Imaging File Types](#_bookmark35)

[Offline Message](#_bookmark38) [Version Compatibility](#_bookmark38)

* [VistA Scratch HFS Directory](#_bookmark43)

1 Patch MD\*1.0\*4 September 2006 Imaging Network Share directory name changed to CP/BGP Transfer Directory.

* 1. Click **System Parameters**, which is displayed under the Clinical Procedures folder. The System Parameters Edit window is displayed. See [Figure 6-10](#_bookmark28).
  2. Enter information in the necessary fields and in the optional fields as needed by your site.



**[1](#_bookmark29)Figure 6-10**

### Allow non-instrument attachments

Select if you want to let users attach files from the network to studies. If selected, the +Files icon displays in the Study window in CP User and lets the user select attachments. Indicates if external attachments (documents) are allowed including when an instrument has not created data.

Be sure to select **Allow non-instrument attachments** if you selected the **Require External Data field** in **CP Manager** for a specific procedure. If you do not select Allow non-instrument attachments, you will not be able to attach files to a procedure.

### Bypass CRC Checking

Select if you want to bypass CRC (Cyclical Redundancy Check) during startup. When a CP application starts up, it can check with the server to be sure that the checksum of the application that is running is the same as the checksum of the application that was distributed. If the checksum values do not match, a message displays stating that the values do not match. Even if values don’t match, you can continue using CP.

The checksum value is associated with the version number of the software. You might want to bypass this check when your site is running CP in test mode. If you are running different versions of the application, then the checksum values will not match.

### Clinical Procedures Home Page

Displays the Clinical Procedures home page and directs the browser to this page when accessed. This parameter is used by the client application in the Help menu when the user selects the option Clinical Procedures on the Web.

**Note**: The MDPOST routine in the KIDS build sets this field during installation. The data in the parameter is predefined. Do not modify this parameter unless the site is performing local modifications to the client software.

### Clinical Procedures On-Line

Must select if you want to use CP User and CP Gateway. If this parameter is not selected, a warning message is displayed. (If a message has been entered into the Offline Message parameter, that message is displayed when the user tries to access CP User.)

This parameter is only effective when the VistA system is functioning and it is useful if you want to restrict access to Clinical Procedures. For example, you can set this field to offline if you are loading a newer version of CP.

### CP/BGP Transfer Directory


Enter the shared directory that is accessed by the Imaging Background Processor (BGP) and CP Gateway. Reports generated from text need to be placed in a location that can be accessed by the BGP. The Network share must not reside physically on the Imaging RAID. You can also use Browse to select the directory. Must be filled in for CP to work properly.

### CRC Values

A site can check that a specific build of the application is running on the client. This level of checking is not mandatory and you can use the Bypass CRC Checking parameter if the site does not want this level of security.

If a site is running more than one version of the application or is testing a new patch, this field can contain multiple entries, ([Figure 6-11](#_bookmark32)). Each entry contains the name of the application with extension (no directory path) followed by a colon ‘:’ and the executable version number ‘#.#.#.#’. Each of these entries contains the CRC value for that particular version of the executable. You can also obtain CRC values for a version of an executable from the About menu or by selecting **CP Manager > Tools > Calculate a File’s CRC Value**.

**Note**: The MDPOST routine in the KIDS build sets this field during installation. The data in the parameter is predefined. Do not modify this parameter unless the site is performing local modifications to the client software



**[1](#_bookmark33)Figure 6-11**

### Calculating a File’s CRC Value

You can calculate a file’s CRC (Cyclical Redundancy Check) value to determine if the file is the exact same file as the one that was distributed. CRC values are recalculated every time an application is compiled.

##### Select Tools > Calculate a file’s CRC Value.

1. Select the file.
2. You can copy the CRC value and paste it into a text file for reference purposes.

### Days to keep instrument data


Enter the number of days (0-365) to save data from auto-instruments, after the data has been associated with a Clinical Procedures study. If the data has not been associated with a study, the data is not purged from the temporary storage area. Enter 0 or leave the field empty if you want the data to be retained forever.

**Note**: CP Gateway purges data daily. This purge only deletes the raw data that comes from the instrument. CP Gateway keeps data for a specified number of days based on the entry in “Days to keep Instrument Data”. Data older than this is purged. The data in Item Value field (#.1) and Item Text field (#.2) of the Upload Item multiple in the CP Results file (#703.1) are purged.

### Imaging File Types

Verifies that a file type submitted by an instrument or user is acceptable and can be sent to the VistA Imaging RAID. The Open a Study option in CP User uses this system parameter to determine if a file is an acceptable file type, ([Figure 6-12](#_bookmark36)).


**Note**: The MDPOST routine in the KIDS build sets this field during installation. The data in the parameter is predefined. Do not modify this parameter unless the site is performing local modifications to the client software



**[1](#_bookmark37)Figure 6-12**

### Offline Message

Enter a message that users see when they try to activate CP User and Clinical Procedures is offline. This message only displays when the Clinical Procedures On-line parameter is not checked. See [Figure 6-13](#_bookmark39).



**[1](#_bookmark40)Figure 6-13**

### Version Compatibility


Displays a list of client versions, identified by their executable name and windows file version, which are compatible with the currently running server version. More than one version of the software may be flagged as compatible for backward compatibility. See [Figure 6-14](#_bookmark41).

To check the client version number:

1 Patch MD\*1.0\*4 September 2006 Imaging Network Share directory name changed to CP/BGP Transfer Directory.

1. Open **Windows Explorer** and locate the Clinical Procedures folder.
2. Right-click CPGateway.exe, or CPUser.exe., or CPManager.exe.
3. Select **Properties**, and then click the **Version** tab. The version number, such as 1.0.0.17, is displayed.
4. Go back to **CP Manager**. Double-click **Clinical Procedures**, and then click **System Parameters**.
5. In the **Version Compatibility** tab, select each version that is compatible with the current server version, ([Figure 6-14](#_bookmark41)).

**Note**: The MDPOST routine in the KIDS build sets this field during installation. The data in the parameter is predefined. Do not modify this parameter unless the site is performing local modifications to the client software



**[1](#_bookmark42)Figure 6-14**

If an executable version is not compatible, the following message is displayed when you try to use a Clinical Procedures application:



**Figure 6-15**

If the application is CP Manager, the user is allowed to continue. If the application is CP User, the user needs to contact IRM because the client needs to be upgraded to the current version.

### VistA Scratch HFS Directory

Clinical Procedures uses the Host File Server (HFS) functionality in the VA Kernel to create reports. VistA broker processes require full read, write, and delete access to this directory. (Check with IRM about this directory.) If this directory is not filled in, CP tries to use the broker environment directory. Must be filled in for CP to work properly.

## [1](#_bookmark45)Step 5 – Exported Kernel XPAR Parameters

### Exported Kernel XPAR Parameters for Patch MD\*1.0\*14

There are four Kernel XAR Parameters exported with patch MD\*1\*14.

* + MD CHECK-IN PROCEDURE LIST
  + MD CLINIC QUICK LIST
  + MD CLINICS WITH MULT PROC
  + MD USE APPT WITH PROCEDURE

A new option called MD AUTO CHECK-IN SETUP was added to setup and implement procedures that will use auto study check-in. Once a procedure is set up to use the auto study check-in functionality in the MD CHECK-IN SETUP option, the software will check-in any existing order requests with the status of “PENDING,” “ACTIVE,” and “SCHEDULED” in the Consult Request Tracking package.

**Note:** If your site uses appointments, schedule them **before** you enter the procedures for auto check-in. If you do not, the patients associated with those appointments will need to be manually checked in.

This option collects the following information:

* 1. Use Appointment with procedure? (Yes/No) (Required) – The default is “NO”, if the site does not schedule procedures before the order is entered. Enter “YES” if the procedure appointment is scheduled before the order is entered and the ordering provider selects the appointment for the procedure during ordering in CPRS.
  2. Procedure (Required)– Enter the CP Definition that will be using the auto study check-in functionality.
  3. Schedule Appointment? (Required) - Enter 0 for None, 1 for Outpatient, 2 for Inpatient, or 3 for Both. This indicates that the site schedules appointments for inpatient, outpatients, both, or none.
  4. Clinic (Optional) – Enter the hospital location(s) that will be used for scheduling the procedure.

**Note:** If no clinic is entered in the setup, CP will use the hospital location defined in the HOSPITAL LOCATION field of the CP Definition file (#702.01) as the location of the visit for the CP study check-in.

The following two pages contain a screen capture of the MD AUTO CHECK-IN SETUP option:

1 Patch MD\*1.0\*14 March 2008 Exported Kernel XPAR Parameters, option, and screen sample added.

Select OPTION NAME: MD AUTO CHECK-IN SETUP Auto Study Check-In Setup Auto Study Check-In Setup

Use Appointment with procedure? NO// ?

Default should be 'N' as most sites do not schedule procedures

before the order is entered. Select 'Y' if the procedure appointment is scheduled before the order is entered and the ordering provider selects the appointment for the procedure.

Enter either 'Y' or 'N'.

Use Appointment with procedure? NO// Procedure: ?

Enter a CP Definition for the procedure to have auto CP study check-in.

Answer with CP DEFINITION NAME

Do you want the entire CP DEFINITION List? N (No)

Procedure: COLONOSCOPY Schedule Appointment?: ?

REQUIRED field for the procedure to have auto CP study check-in. Enter a "^" will exit completely.

Enter 0 if you do not schedule appointments.

1. if you only schedule appointments for outpatients.
2. if you only schedule appointments for inpatients.
3. if you schedule appointments for both 1 and 2. Select one of the following:
   1. None
   2. Outpatient
   3. Inpatient
   4. Both

Schedule Appointment?: Both Clinic: ?

Only required, if appointments are scheduled for the procedure. Enter the clinic used for scheduling the procedure.

Answer with HOSPITAL LOCATION NAME, or ABBREVIATION, or TEAM

Do you want the entire 112-Entry HOSPITAL LOCATION List? N (No) Clinic: GI LAB PIPER,ALPHA

Enter another clinic for the same procedure? NO// ?

Enter either 'Y' or 'N', if you want to assign more than one clinic. Enter another clinic for the same procedure? NO//YES

Clinic: TEST

1. TEST/PROSTHETICS OBRIEN,FRANCES U
2. TEST1
3. TEST1234
4. TEST3232

CHOOSE 1-4: 2 TEST1

Enter another clinic for the same procedure? NO// Procedure: ?

Enter a CP Definition for the procedure to have auto CP study check-in.

COLONOSCOPY

Answer with CP DEFINITION NAME

Do you want the entire CP DEFINITION List? N (No)

Procedure: EKG, ROUTINE (12 LEADS)

Schedule Appointment?: 0 None

Procedure:

>

**[1](#_bookmark47) Exported Kernel XPAR Parameters for Patch MD\*1.0\*6** There are four Kernel XPAR Parameters exported with Patch MD\*1.0\*6. PARAMETER DEFINITION:

* MD APPOINT END DATE
* MD APPOINT START DATE
* MD COMPL PROC DISPLAY DAYS
* MD DAYS TO RETAIN COM STUDY

The users can edit the parameters using the Edit Parameter Values option, [XPAR EDIT PARAMETER].

The following is a screen capture of the parameter usage:

Select PARAMETER DEFINITION NAME: MD APPOINT END DATE End Date for Encounter Appointments

---- Setting MD APPOINT END DATE for System: REDACTED ----

Days: ?

Enter a number from 0 to 365. Days: ??

Enter a number from 0 to 365 for the number of days that will be used to add to today as the end date range of the Encounter Appointments. If no value is entered, the default value used will be 0.

Edit Parameter Values

--- Edit Parameter Values ---

Select PARAMETER DEFINITION NAME: MD APPOINT START DATE Start Date for Encounter Appointments

--- Setting MD APPOINT START DATE for System: REDACTED ---

Days: ?

Enter a number from 0 to 365. Days: ??

Enter a number from 0 to 365 for the number of days that will be used to subtract from today as the start date range of the Encounter Appointments. If no value is entered, the default value used

will be 200.

Days: 365

Edit Parameter Values

D ^XUP

Setting up programmer environment Terminal Type set to: C-VT100

You have 2983 new messages.

Select OPTION NAME: XPAR EDIT PARAMETER

1 Patch MD\*1.0\*6 May 2008 Exported Kernel XPAR Parameters and screen sample added.

Days: 2

Select PARAMETER DEFINITION NAME: MD COMPL PROC DISPLAY DAYS Completed Proc Display Days Setting MD COMPL PROC DISPLAY DAYS for System: REDACTED

Days: ?

Enter the number of days from 1 to 365. Days: ??

The number of days the completed procedure requests will be displayed in the CP Check-in screen.

Days: 365

Select PARAMETER DEFINITION NAME: MD DAYS TO RETAIN COM STUDY Days to Retain Completed Study

Setting MD DAYS TO RETAIN COM STUDY for System: REDACTED

Days: ?

Enter the number of days from 1 to 365. Days: ??

The number of days after check-in date/time to display the study that has been complete in the CPUser application. Studies that have procedures with multiple or cumulative results are NOT included.

Cumulative and multiple results studies will have a default value of 365.

Days:

Select PARAMETER DEFINITION NAME:

# 11. Setting Up HL7 Parameters

[1](#_bookmark50)This section describes how to set up the HL7 parameters including configuration instructions, file settings, and technical issues. The tasks in this chapter require a working knowledge of the VistA HL7 application.

Topics discussed in this chapter are:

* [Configuration Instructions Information](#_bookmark49)
  + [IP Addresses and Ports](#_bookmark52)
* [Setting Up a New HL7 Single Listener for High-Volume Devices](#_bookmark53)
  + [Creating a Logical Link](#_bookmark55)
  + [Creating a Device Protocol Client](#_bookmark57)
  + [Activating the Logical Links](#_bookmark59)
  + [Adding a Device Client as a Server Subscriber](#_bookmark60)
* [Using Port 5000](#_bookmark62)
  + [Benefits of Using a Single Port Listener](#_bookmark63)
  + [Setting Up Port 5000](#_bookmark64)
* [File Settings](#_bookmark65)
* [Technical Issues](#_bookmark68)

## Configuration Instructions Information

You can follow the steps described in this section to configure the HL7 application.

MCAR INST and MCAR OUT are automatically created during the KIDS installation. MCAR INST is used for all devices that send results information from the device to VistA and CP. Since all devices can use the same link, you only need one entry in the HL Logical Link (870) file.

However, you need to establish an MCAR OUT entry for each bi-directional device that receives information from VistA and CP. Each entry needs its own IP and port number, which agree with the device configuration. (Use the MCAR OUT sample provided in the HL Logical Link file. Set up the individual links for each bi-directional device. ) All outbound links are non-persistent.

Most devices are able to use a non-persistent connection to VistA. A **persistent** connection is a connection that is established by the medical device and is kept connected to VistA even after the device has transmitted it HL7 message. A **non-persistent** connection is a connection that is established by the medical device to VistA and is disconnected once the HL7 message has been sent. Devices can share the same HL Logical link to VistA, if they are non-persistent. If the device is persistent then it must have its own HL Logical Link to VistA (Example: its own inbound and outbound links.)

1 Patch MD\*1.0\*14 March 2008 Chapter revised to provide clarity.

### IP Addresses and Ports

You need to set up IP addresses and ports for the medical devices at your facility.

An IP address consists of a string of four numbers each ranging in value from 0 to 255. Here is an example of an IP address: 10.23.55.201. When a new device is installed, be careful when you assign IP addresses to the medical devices. It’s recommended that you set aside a block of IP addresses specifically for the medical devices. The range of numbers chosen is up to the facility, but make sure that there is a large enough range to allow for some growth. For example, IP addresses 10.23.55.201 through 10.23.55.225 could be blocked and used. In this way, the IRM staff can track down any possible problems that may be related to the medical device by looking at the IP address.

A port is the location on a medical device where you send and receive data. Some ports have predefined functions. For example, Port 80 is set up for the Web Server. Some vendors have predefined ports that they may want you to use. For example, Sensormedics recommends using Port 20000 for the VMAX. Others may only allow a limited range. Consult the device manual to determine which ports you can use.

A Startup Node defines the system on which you want the link to start.

## Setting Up a New HL7 Single Listener for High-Volume Devices

Most medical devices send results to VistA using nonpersistent connections to the same port. Each device connects to the port just long enough to send results to VistA, then releases the port so that other devices may connect to it.

However, if you use a high-volume device (i.e., something that sends about 200 or more messages back and forth per day, such as MUSE or a hemodialysis device) that sends a lot of data all the time, we recommend that you give it its own port instead of sharing a port with other devices. This is because high-volume devices send so much data that they can tie up the port for a long time, preventing other devices (e.g., Olympus or Sensormedics) from using it.

Setting up a new HL7 listener involves four steps (which are described in more detail below):

1. [Creating a Logical Link](#_bookmark55)
2. [Creating a Device Protocol Client](#_bookmark57)
3. [Activating the Logical Links](#_bookmark59)
4. [Adding a Device Client as a Server Subscriber](#_bookmark60)

This document also contains information on [Using Port 5000](#_bookmark62) what it is and when to use it).

**Note:** Although you can name your new logical links and device protocols anything you want, keep the names name spaced and descriptive since the names are similar and it can be easy to confuse them.

|  |
| --- |
| **Creating a Logical Link**  A logical link is an inbound or an outbound instrument data port from and to the medical device. It’s a listener waiting for data to come across. The first logical link (MCAR INST) is already created by default. To create a new HL7 single listener logical link for your device, you need to create a new logical link or edit an existing one. Each bi-directional device that receives information from VistA and CP needs its own outbound link set up. All non-persistent devices can share the same inbound logical link, but persistent devices each need their own inbound logical links.   1. Decide which port to use. The facility, along with IRM, determines which port to use. This is the port used by the device to send data to the VistA listener. You can, for example, use port 1026 for Hemodialysis results and port 1027 for Sensormedics results. Do not use port 5000 for this type of setup. (See below for more information on port 5000.) 2. From the Systems Manager Menu, choose HL Main Menu (**HL**) > Interface Developer Options (**IN**) > Link Edit (**EL**). 3. At the Select HL Logical Link Note prompt, enter the **name of the new logical link** for your device. Name your new inbound logical link something like MCAR2 INST. The next one (if you use more than one high-volume device) can be called MCAR3 INST, etc. For an outbound logical link, the following naming convention is suggested: MCAR xxx, where xxx is the first three characters of the device or vendor name. (For example, an outbound link for an Olympus device could be named MCAR OLY.) 4. Type **yes** when asked if you are adding ‘MCAR2 INST’ as a new HL LOGICAL LINK. The HL7 LOGICAL LINK screen displays. |
| HL7 LOGICAL LINK  NODE: MCAR2 INST INSTITUTION:  MAILMAN DOMAIN:  AUTOSTART: **Enabled**  QUEUE SIZE: 100  LLP TYPE: TCP **<RET>**  DNS DOMAIN: |
| **Note:** When this screen first displays for a new logical link, only the NODE and QUEUE SIZE fields will already contain values. The NODE field will display the logical link name you just created, and the QUEUE SIZE field will default to 10.   1. Type **Enabled** in the AutoStart field. 2. Change the QUEUE SIZE value to 100. (Optional) |

1. Enter **TCP** in the LLP TYPE field, then press **[Enter]** to display the HL7 LOGICAL LINK screen (see following figure).

HL7 LOGICAL LINK

┌─────────────────────TCP LOWER LEVEL PARAMETERS──────────────────────────┐

|  |  |
| --- | --- |
| │ MCAR3 INST | │ |
| │ | │ |
| │ TCP/IP SERVICE TYPE: **SINGLE LISTENER** | │ |
| │ TCP/IP ADDRESS: | │ |
| │ TCP/IP PORT: 1026 | │ |
| │ TCP/IP PORT (OPTIMIZED): | │ |
| │ | │ |
| │ ACK TIMEOUT: 60 RE-TRANSMISION ATTEMPTS: 3 | │ |
| │ READ TIMEOUT: EXCEED RE-TRANSMIT ACTION: ignore | │ |
| │ BLOCK SIZE: SAY HELO: | │ |
| │ | │ |
| │STARTUP NODE: DEV:ISC4A2 PERSISTENT: NO | │ |
| │ RETENTION: UNI-DIRECTIONAL WAIT: | │ |

└─────────────────────────────────────────────────────────────────────────┘

1. Set TCP/IP SERVICE TYPE to **SINGLE LISTENER**. If this link is an outbound link then the TCP/IP SERVICE TYPE is **CLIENT (SENDER)**.
2. In the **TCP/IP ADDRESS** field, if the link you are creating is an outbound link to a device, you will need to enter the TCP/IP address of that device. For the inbound link, no address is needed.
3. In the **TCP/IP PORT** field, enter the port number you decided to use (in step 1).
4. Optionally set ACK TIMEOUT to **60**.
5. Press **[Tab]** to optionally set RE-TRANSMISION ATTEMPTS to **3**.
6. Optionally set EXCEED RE-TRANSMIT ACTION to **ignore**.
7. Enter the appropriate **STARTUP NODE**.
8. Set the PERSISTENT field to **NO**.
9. **[Tab]** down to the COMMAND prompt, then select **Close**. You return to the HL7 LOGICAL LINK screen.
10. **[Tab]** down to the COMMAND prompt, then select **Save**.
11. At the COMMAND prompt, select **Exit**.
12. The new link is useless until you assign protocols to it. Proceed to the next section to create a client protocol.

Below is an example of an outbound link.

┌──────────────────────TCP LOWER LEVEL PARAMETERS─────────────────────────┐

│ MCAR OLY │

│ │

│ TCP/IP SERVICE TYPE: CLIENT (SENDER) │

│ TCP/IP ADDRESS: 10.3.17.141 │

│ TCP/IP PORT: 9027 │

│ TCP/IP PORT (OPTIMIZED): │

│ │

│ ACK TIMEOUT: 60 RE-TRANSMISION ATTEMPTS: 3 │

│ READ TIMEOUT: 60 EXCEED RE-TRANSMIT ACTION: ignore │

│ BLOCK SIZE: SAY HELO: │

│ DIRECT CONNECT OPEN TIMEOUT: │

│STARTUP NODE: DEV:DEVISC4A1 PERSISTENT: NO │

│ RETENTION: UNI-DIRECTIONAL WAIT: │

└─────────────────────────────────────────────────────────────────────────┘

### Creating a Device Protocol Client

You have to create a protocol for every inbound listener to VistA.

To create a protocol client from for your new logical link using a **copy**, follow these steps:

1. Look at the protocol in 101 or use developer tools. Copy MCAR DEVICE CLIENT to make a new device client. Name it something like MCAR2 DEVICE CLIENT.
2. Change the entry in the Logical Link field to match the new logical link. For example, if you just created a logical link named MCAR2 INST, change what’s in the Logical Link field from MCAR INST to MCAR2 INST. All other fields should match what was originally in MCAR DEVICE CLIENT.
3. Proceed to the next section to make the new device protocol a subscriber to the device server.

To create a **new** protocol client for your new logical link, do the following:

1. From the Systems Manager Menu, choose HL Main Menu (**HL**) > Interface Developer Options (**IN**) > Protocol Edit (**EP**).
2. At the Select PROTOCOL NAME prompt, enter the name of the new device client for your device. Name your new device client something like MCAR2 Device Client or MCAR2 MUSE (depending on the device name).
3. Type **yes** (or simply type **y**) when asked if you are adding ‘MCAR2 Device Client’ as a new PROTOCOL.
4. Enter **Instrument Device Client** in the PROTOCOL ITEM TEXT field.
5. Enter an appropriate identifier in the PROTOCOL IDENTIFIER field. The HL7 INTERFACE SETUP screen displays.
6. **[Tab]** down to the TYPE field and enter **subscriber**, then press **[Enter]** to display PAGE 2 OF 2.

HL7 SUBSCRIBER PAGE 2 OF 2

MCAR2 Device Client

RECEIVING APPLICATION: MCAR INST

RESPONSE MESSAGE TYPE: ACK EVENT TYPE: R01 SENDING FACILITY REQUIRED?: NO RECEIVING FACILITY REQUIRED?: NO

SECURITY REQUIRED?:

LOGICAL LINK: MCAR2 INST PROCESSING RTN: D ^MDHL7A

ROUTING LOGIC:

1. Type **MCAR-INST** in the RECEIVING APPLICATION field, then enter the following entries:
2. RESPONSE MESSAGE TYPE = **ACK**
3. EVENT TYPE = **R01**
4. SENDING FACILITY REQUIRED = **NO**
5. RECEIVING FACILITY REQUIRED = **NO**
6. LOGICAL LINK = **MCAR2 INST** (use the appropriate name)
7. PROCESSING RTN = **D ^ MDHL7A** (use the appropriate routine)

**Note:** The processing routine is the MUMPS routine that VistA uses to process the message received from the logical link.

1. **[Tab]** down to the COMMAND prompt, then select **Save**.
2. At the COMMAND prompt, select **Exit**.
3. Proceed to the next section to make the new device protocol a subscriber to the device server.

### Activating the Logical Links

Next, the links need to be activated. (The steps below assume that the original logical link has never been activated. If MCAR INST is already active, skip to step 4.)

1. Choose HL Main Menu (**HL**) > Filer and Link Management Options (**FI**) >Start/Stop Links (**SL**).
2. Activate the first logical link: Select HL LOGICAL LINK NODE: **MCAR INST**
3. Select **B** for Background. (B is the default, so just press **[Enter]**.
4. Activate the next logical link: Select HL LOGICAL LINK NODE: (in this example it is

##### MCAR2 INST)

1. Select **B** for Background. (B is the default, so just press **[Enter]**.
2. If you have more logical links to activate, repeat steps 4-5.
3. If you haven’t done this already, use the CP Manager application to configure the device you are using. Refer to [Editing an Automated Instrument](#_bookmark10), p. 6-3.
4. Proceed to the next section to make the new device protocol a subscriber to the device server.

### Adding a Device Client as a Server Subscriber

Next you have to make the newly-created protocols subscribers to MCAR DEVICE SERVER. Every client must be a subscriber to a server. That controls the outbound message to a medical device when you reply to it.

Go into MCAR DEVICE SERVER (under the protocol file or using the Interface Developer Option) and make sure that the new MCAR2 DEVICE CLIENT is a subscriber to it. Detailed steps follow:

1. At the Select Systems Manager Menu, select **HL** for the HL7 Main Menu.
2. At the Select HL7 Main Menu, select **IN** for Interface Developer Options.
3. At Select Interface Developer Options, select **EP** for Protocol Edit.
4. At the Select PROTOCOL NAME prompt, select **MCAR Device Server**. (If your site uses a different server name, select the appropriate name. You can display a list of available options, if necessary.)
5. Press **[Enter]** at the TYPE prompt to go to PAGE 2 OF 2: the HL7 EVENT DRIVER screen. (Figure follows.)

HL7 EVENT DRIVER PAGE 2 OF 2

MCAR Device Server

SENDING APPLICATION: INST-MCAR

TRANSACTION MESSAGE TYPE: ORU EVENT TYPE: R01 MESSAGE STRUCTURE:

PROCESSING ID: P VERSION ID: 2.3 ACCEPT ACK CODE: APPLICATION ACK TYPE:

RESPONSE PROCESSING RTN:

SUBSCRIBERS

MCAR Device Client MCAR Device Client2

1. To add the new protocol as a subscriber, **[Down Arrow]** or **[Tab]** down to the line below MCAR Device Client and **enter the name of the new subscriber** (e.g., MCAR2 Device Client). The HL7 screen displays.
2. Verify that the entries are correct, then **[Down Arrow]** down to the COMMAND line and select **Close**. You return to the MCAR Device Server screen.
3. Repeat steps 6-7 if you need to add more subscribers.
4. **[Down Arrow]** down to the COMMAND line and select Save.
5. In the COMMAND line, select **Exit**.

## [1](#_bookmark66)Using Port 5000

Port 5000 is a Multi-Port Listener. The only reason to use the multiport listener is if your inbound port doesn’t work correctly because Cache is not handling ports correctly.

If Cache is handling ports correctly, then you should let Cache handle them. Use the individually shared ports for your devices rather than using the Multi-Port Listener.

If you’re at a facility that has listener problems under Cache, then use port 5000. Port 5000 is handled by VMS, not Cache.

Most sites allocate 25 ports to port 5000, but more can be allocated, if necessary.

### Benefits of Using a Single Port Listener

A single port is easier to monitor and debug. It’s easy to determine if the problem is caused by the link or something else.

If you set up another Multi-Port Listener, you have to set it up in VMS. You’ll have to do that through UCX, which is a lot of work and beyond the scope of this document,

### Setting Up Port 5000

1. Edit MCAR DEVICE CLIENT so the logical link points to VAxxx (where xxx is an abbreviation for the hospital).
2. Make sure all CP Medical devices send to port 5000.
3. You don’t need to set up an additional MCAR INST (logical link) because you’re using an existing logical link which is VA**xxx**, where xxx is an abbreviation for the hospital (e.g., VAHIN for Hines).
4. Make it an MCAR DEVICE server subscriber.

## File Settings

The parameter settings for the HL7 Application Parameter file, HL Logical Link file, and the Protocol file are automatically set during the CP installation. They are listed here for reference. Fields that have bolded field names and bolded field entries must be set exactly as they appear in these examples.

##### HL7 Application Parameter (#771) file

This file contains a list of VistA applications that are capable of sending and receiving HL7 transmissions.

NAME: **MCAR-INST ACTIVE/INACTIVE: ACTIVE FACILITY NAME: VISTA MAIL GROUP: POSTMASTER**

**COUNTRY CODE: US HL7 ENCODING CHARACTERS: ^~\& HL7 FIELD SEPARATOR: |**

NAME: **INST-MCAR ACTIVE/INACTIVE: ACTIVE**

1 Patch MD\*1.0\*9 November 2007 Using Port 5000 with CACHE.

**COUNTRY CODE: US HL7 ENCODING CHARACTERS: ^~\& HL7 FIELD SEPARATOR: |**

##### HL Logical Link (#870) file

This file stores parameters that govern the behavior of the Logical Links and also stores information that drives the SYSTEMS LINK MONITOR display option.

NODE: **MCAR INST LLP TYPE: TCP**

QUEUE SIZE: 100 RE-TRANSMISSION ATTEMPTS: 3

ACK TIMEOUT: 60 EXCEED RE-TRANSMIT ACTION: ignore

TCP/IP PORT: 1026 **TCP/IP SERVICE TYPE: SINGLE LISTENER PERSISTENT: NO**

MCAR OUT provides an example of field entries for bi-directional instruments for outbound links to medical devices. The fields that have bolded field names and bolded field entries must be set exactly as they appear in this example. The other bolded fields must be edited to match your device specific requirements. For example, Device Type must be Non-Persistent Client. Non- bolded fields may not have a value depending on the state of the system.

**NODE**: MCAR OUT **LLP TYPE: TCP**

**DEVICE TYPE: Non-Persistent Client** STATE: Shutdown

**AUTOSTART: Enabled** TIME STOPPED: JAN 16, 2003@14:30:15

**SHUTDOWN LLP ?: YES EXCEED RE-TRANSMIT ACTION: ignore**

**RE-TRANSMISSION ATTEMPTS: 3 TCP/IP PORT:** 1028

**ACK TIMEOUT: 60 PERSISTENT: NO**

**TCP/IP ADDRESS:** 10.3.17.202 **STARTUP NODE:** DEV:ISC4A2

**TCP/IP SERVICE TYPE: CLIENT (SENDER)**

**Note:** When you need to create additional HL7 links for new devices, name the link in the following format:

* If you need to create more than one inbound link (MCAR INST), name the new links “MCAR”, followed by a number (1,2,3), a space, and then “INST”.

Example: MCAR2 INST

* Name outbound links “MCAR”, followed by a number (1,2,3), a space, and then a name for the device.

Example: MCAR2 SMC

[See “Configuration Instructions Information](#_bookmark49)” for information on setting the TCP/IP address and port and the Startup Mode.

##### Protocol (#101) file:

This file contains the protocols for processing HL7 messages.

NAME: **MCAR Device Client ITEM TEXT: Instrument Device Client TYPE: subscriber** CREATOR: CPUSER, FIVE

**PACKAGE: MEDICINE**

DESCRIPTION: Subscriber protocol for sending data to VISTA from clinical instruments.

TIMESTAMP: 57540,31165 **RECEIVING APPLICATION: MCAR INST TRANSACTION MESSAGE TYPE: ORU EVENT TYPE: R01**

**PROCESSING ID: P LOGICAL LINK: MCAR INST**

**\* VERSION ID: 2.3 RESPONSE MESSAGE TYPE: ACK PROCESSING ROUTINE: D ^MDHL7A SENDING FACILITY REQUIRED?: NO RECEIVING FACILITY REQUIRED?: NO**

NAME: **MCAR Device Server ITEM TEXT: Instrument HL7 Event Driver TYPE: event driver** CREATOR: CPUSER, FIVE

**PACKAGE: MEDICINE**

DESCRIPTION: This protocol is used by the HL7 package to send results to VISTA from various clinical instrumentation.

TIMESTAMP: 57631,55707 **SENDING APPLICATION: INST-MCAR TRANSACTION MESSAGE TYPE: ORU EVENT TYPE: R01**

**PROCESSING ID: P \* VERSION ID: 2.3**

**SENDING FACILITY REQUIRED?: NO RECEIVING FACILITY REQUIRED?: NO**

**SUBSCRIBERS: MCAR Device Client**

NAME: **MCAR ORM CLIENT TYPE: subscriber**

CREATOR: CPUSER, SIX **RECEIVING APPLICATION: INST-MCAR**

**EVENT TYPE: O02 RESPONSE MESSAGE TYPE: ORR SENDING FACILITY REQUIRED?: NO RECEIVING FACILITY REQUIRED?: NO SECURITY REQUIRED?: NO ROUTING LOGIC: Q**

NAME: **MCAR ORM SERVER**

**ITEM TEXT: Clinical Procedures ORM Protocol Server**

**TYPE: event driver** CREATOR: CPUSER, SIX

TIMESTAMP: 59276,54156 **SENDING APPLICATION: MCAR-INST TRANSACTION MESSAGE TYPE: ORM EVENT TYPE: O01**

**VERSION ID: 2.3 SUBSCRIBERS: MCAR ORM CLIENT**

**\*Note:** Check vendor documentation for instructions on verifying the Version ID.


## Technical Issues

For all sites:

To avoid error messages because of a missing or invalid 'Event Protocol', 'Invalid Processing Code', or 'Invalid Application Code', make sure that all settings (except TCP/IP PORT and TCP/IP ADDRESS, in the HL Logical Link (#870) file, which are site specific) are the same as the file settings listed previously in this chapter.

Be sure that the VERSION ID parameters in the Protocol (#101) file are set to the same HL7 Version that is being sent by the vendor instrument. The ITEM and SUBSCRIBERS fields in


the Device Server entry in the Protocol (#101) file MUST be the same as the Device Client name.



[**1**](#_bookmark69)**Figure 13-3**

1. An error status is displayed for the study and the **Update Study Status** selection is unavailable. You must have the MD GUI MANAGER key, and then you can go to **File > Update Study Status** to review the problem.

The message in the following figure indicates that a **Notification Mailgroup** has not been assigned or the **Medical Device** is not **Active**.



**Figure 13-4**


1 Patch MD\*1.0\*6 May 2008 Added Processing Application field to image.

* 1. Open **CP Manager**.
  2. Select the instrument.
  3. Check that the **Notification Mail Group** has an entry and that the **Active** checkbox is selected.
  4. Open **CP User**. Choose **File > Update Study Status**.
  5. If the device is bi-directional, delete the study that was checked in and check-in a new study with the same procedure request to get the HL7 message transmitted to the medical device. If the device is uni-directional, check the **Ready to Complete** status, and click **OK**.

1. If a study remains in **Pending Instrument Data** status and it is a bi-directional medical device, check to see if **Auto Submit To VistA Imaging** field is selected. .
   1. Open **CP Manager.**
   2. Expand the Procedures folder, and then select the procedure
   3. Check that **Auto Submit to VistA Imaging** is selected.

For the current study, you still need to manually submit the result. For future studies, the result will be automatically submitted.

1. The following two errors indicate that a TIU document Title has not been assigned to the CP procedure. The first error message is from CP during image submission if a TIU document has not been assigned to the CP Definition.



**Figure 13-5**

This second error screen is the Update Study Status screen from CP User. The first message is a CP warning. The second message is a warning from TIU that there is no TIU document.

**Document Class** Document Classes are categories that group documents (Titles) with similar characteristics together. For example, Cardiology notes might be a Document Class, with Echo notes, ECG notes, etc. as Titles under it. Or maybe the Document Class would be Endoscopy Notes, with Colonoscopy notes, etc. under that Document Class.

**Document Definition** Document Definition is a subset of TIU that provides the building blocks for TIU, by organizing the elements of documents into a hierarchy structure. This structure allows documents (Titles) to inherit characteristics (such as signature requirements and print characteristics) of the higher levels, Class and Document Class. It also allows the creation and use of boilerplate text and embedded objects.

**Edit** Used to change/modify data typically stored in a file.

**Field** A data element in a file.

**File** The M construct in which data is stored for retrieval at a later time. A computer record of related information.

**File Manager or FileMan** Within this manual, FileManager or FileMan is a reference to VA FileMan. FileMan is a set of M routines used to enter, edit, print, and sort/search related data in a file, a database.

**File Server** A machine where shared software is stored.

**Gateway** The software that performs background processing for Clinical Procedures.

**GUI** Graphical User Interface – a Windows interface that uses pull-down menus, icons, pointer devices, and other metaphor-type elements that make a computer program, easier to use and that allows multi-processing (more than one window or process available at once).

**Interpreter** Interpreter is a user role exported with USR\*1\*19 to support the Clinical Procedures Class. The role of the Interpreter is to interpret the findings or results of a clinical procedure. Users who are authorized to interpret the results of a clinical procedure are sent a notification when an instrument report and/or images for a CP request are available for interpretation. Business rules are used to determine what actions an interpreter can perform on a document of a specified class, but the interpreter themselves are defined by the Consults application. These individuals are ‘clinical update users’ for a given consult service.

**IRMS** Information Resource Management Service.

**Kernel** A set of software utilities. These utilities provide data processing support for the application packages developed within the VA. They are also tools used in configuring the local computer site to meet the particular needs of the hospital. The components of this

operating system include: MenuMan, TaskMan, Device Handler, Log-on/Security, and other specialized routines.

**M** Formerly known as MUMPS or the Massachusetts (General Hospital) Utility Multi- Programming System. This is the programming language used to write all VistA applications.

**Menu** A set of options or functions available to users for editing, formatting, generating reports, etc.

**Modality** Another name for a medical instrument.

**Module** A component of a software application that covers a single topic or a small section of a broad topic.

**Namespace** A naming convention followed in the VA to identify various applications and to avoid duplication. It is used as a prefix for all routines and globals used by the application.

**Network Server Share** A machine that is located on the network where shared files are stored.

**Notebook** This term refers to a GUI screen containing several tabs or pages.

**Option** A functionality that is invoked by the user. The information defined in the option is used to drive the menu system. Options are created, associated with others on menus, or given entry/exit actions.

**Package** Otherwise known as an application.

**Page** This term refers to a tab on a GUI screen or notebook.

**Password** A protected word or string of characters that identifies or authenticates a user, a specific resource, or an access type (synonymous with Verify Code).

**Persistent Connection** A connection that is established by the medical device and is kept connected to VistA even after the device has transmitted it HL7 message.

**Non-persistent Connection** A connection that is established by the medical device to VistA and is disconnected once the HL7 message has been sent.

**Pointer** A special data type of VA FileMan that takes its value from another file. This is a method of joining files together and avoiding duplication of information.

**Procedure Request** Any procedure (EKG, Stress Test, etc.) which may be ordered from another service/specialty without first requiring formal consultation.

**Queuing** The scheduling of a process/task to occur at a later time. Queuing is normally done if a task uses up a lot of computer resources.

**Result** A consequence of an order. Refers to evaluation or status results. When you use the Complete Request (CT) action on a consult or request, you are transferred to TIU to enter the results.

**Security Key** A function which unlocks specific options and makes them accessible to an authorized user.

**Sensitive Information** Any information which requires a degree of protection and which should be made available only to authorized users.

**Site Configurable** A term used to refer to features in the system that can be modified to meet the needs of each site.

**Software** A generic term referring to a related set of computer programs.

**Status Symbols** Codes used in order entry and Consults displays to designate the status of the order.

**Study** See CP Study.

**Task Manager or TaskMan** A part of Kernel which allows programs or functions to begin at specified times or when devices become available. See Queuing.

**Title** Titles are definitions for documents. They store the behavior of the documents which use them.

**TIU** Text Integration Utilities.

**User** A person who enters and/or retrieves data in a system.

**User Class** User Classes are the basic components of the User Class hierarchy of ASU (Authorization/Subscription Utility) which allows sites to designate who is authorized to do what to documents or other clinical entities.

**User Role** User Role identifies the role of the user with respect to the document in question, such as Author/Dictator, Expected Signer, Expected Cosigner, Attending Physician, etc..

**Verify Code** A unique security code which serves as a second level of security access. Use of this code is site specific; sometimes used interchangeably with a password.

**VistA** Veterans Health Information Systems and Technology Architecture.

# Appendix B – Exported Procedures List

These exported procedures are contained in the MDPOST routine. When the INIT^MDPOST routine is run, these entries are added to your CP Definition (#702.01) file:

ABD PARACENTESIS: FOLLOWUP ABD PARACENTESIS: INITIAL ABLATION OF AV NODE FUNCTION AICD INTER/CONDITION

AIRWAY RESISTANCE ANO BIOPSY

ANO CONTROL BLEEDING

ANO DIAGNOSTIC (BRUSHINGS) ANO HOT BIOPSY(IES)

ANO SINGLE TUMOR (HOT/BICAP) ANOSCOPY

ARRHYTHMIA INDUCTION BY PACING ARTERIAL BLOOD GASES

ARTERIAL CANNULATION ARTERIAL PUNCTURE ARTHROC.ASPIR.INJ.INT.JT.BUR ARTHROC.ASPIR.INJ.MAJ.JT.BUR ARTHROCENT.ASPIR.INJ.SM.JT.BUR ASPIRATION

BIOPSY

BIOPSY LUNG, PERCUTANEOUS NDL BIOPSY, PLEURA

BONE MARROW

BONE MARROW INTERPRETATION BRONC DIAGNOSTIC W/BAL BRONC W/BRONC WASHING BRONC W/TRANSBRONC LUNG BX BRONCHIAL BRUSH

BRONCHOSCOPY W/BRONCH BIOPSY BRONCHOSCOPY W/WANG NEEDLE BRONCHOSCOPY, LASER BRONCHOSCOPY, STENT PLACEMENT BRONCHOSCOPY, THERAPEUTIC BRONCOSCOPY/FB REMOVAL

C&P EXAM

CARDIAC CATHETERIZATION CARDIAC REHAB W/O ECG MON CARDIAC REHAB/W ECG MON

CARDIOPULMONARY REHABILITATION CARDIOVERSION, ELECTIVE

CENTRAL VENOUS CANNULATION CHEMOTHERAPY

COL ABL (OTHR THAN SNARE/BI) COL BIOPSY

COL CONTROL HEM.

COL DIAGNOSTIC (BRUSHINGS) COL HOT BIOPSY(IES)

COL REMOVAL FB COL SNARE COLONOSCOPY

COMPREHENSIVE EP EVALUATION CPAP/BIPAP VENTILATION DIALYSIS PROCEDURES, HEMO DIALYSIS TRAINING/COMPLETE DIFFUSION

DILUTION STUDIES FOR CO MEAS ECG

ECG (EKG), RHYTHM STRIP ECG 12 LEAD

ECG 24 HOUR HOLTER MONITOR ECG MONITORING

ECG WITH INTERPRETATION ECG, EVENT RECORDER ECG, RHYTHM TRACING ECG, SIGNAL AVERAGE ECHO

ECHO TRANSESOPHOGEAL SINGLE PL ECHO, 2D M-MODE

ECHO, DOPPLER COLOR FLOW ECHO, DOPPLER, COMPLETE ECHO, TRANSESOPHOGEAL

ECHO, TRANSESOPHOGEAL BIPLANE ECHO, TRANSTHORACIC

EGD

EGD ABL (OTH THAN SNARE/BI) EGD BAND LIGATION

EGD BIOPSY

EGD DIAGNOSTIC (BRUSHINGS) EGD DILATION BALLOON

EGD DILATION WIRE EGD FOREIGN BODY

EGD HOT BIOPSY(IES) / BICAP EGD INJECTION / SCLEROSIS EGD SNARE/SINGLE

EGD TUBE/STENT

EKG, ROUTINE (12 LEADS)

ENDO OF BOWEL POUCH W/ BIOPSY ENDOMYOCARDIAL BIOPSY ENDOSCOPIC ULTRASOUND ENDOSCOPIC ULTRASOUND, BIOPSY ENDOSCOPY OF BOWEL POUCH ENDOTRACHEAL INTUBATION ENTEROSCOPY

EP EVAL OF CARDIO/DEFIB LEADS EP EVAL OF CARDIOVERTER/DEFIB EP EVAL W/ ARRHYTHMIA INDUCT EP EVAL W/ L ATRIAL RECORD

EP EVAL W/ L VENTRIC RECORD EP FOLLOWUP STUDY W/PACING EP STUDY

EPICARDIAL/ENDOCARDIAL MAPPING ERCP

ERCP ABL (OTHR THAN SN/BI) ERCP BALLOON DILATION ERCP BIOPSY

ERCP DEST STONES

ERCP DIAGNOSTIC (BRUSHINGS) ERCP DRAIN, TUBE

ERCP INSERTION OF TUBE/STENT ERCP PRESSURE OF ODDI

ERCP REM STONES

ERCP RMV FB OR CHG OF TUBE ERCP SPHINCTEROTOMY

ES ABLATION (OTHER) ES BAND LIGATION

ES BIOPSY

ES CONTROL BLEEDING

ES DIAGNOSTIC ENDO (BRUSHINGS) ES DILATION (BALLOON)

ES DILATION (WIRE) ES HOT BIOPSY(IES)

ES INJECTION / SCLEROSIS ES INSERTION TUBE/STENT ES REMOVAL FB

ES SNARE

ESOPHAGEAL DILATION ESOPHAGEAL MOTILITY STUDY ESOPHAGEAL RECORDING ESOPHAGUS

ETT

ETT W/ O2 CONSUMPTION ETT W/ THALLIUM SCAN

EXAM,SYNOVIAL FLUID CRYSTALS EXCERCISE CHALLENGE

FINE NEEDLE ASPIRATION FLEX SIG

FLOW VOLUME LOOP FLX ABLATION (OTHER) FLX BIOPSY

FLX CONTROL HEM.

FLX DECOMPRESS VOLVULUS FLX DIAGNOSTIC (BRUSHINGS) FLX HOT BIOPSY(IES)

FLX REMOVAL FB FLX SNARE

FRC

FT CHANGE OF G TUBE

FT EGD FOR PEG PLACEMENT

FT PERC PLACEMENT OF G TUBE FT REPOS TUBE THRU DUODENUM FT SM INT ENDO CONV G-J TUBE FT SM INT ENDO J TUBE PLACE HEART RATE VAR. ANALYSIS HEMODIALYSIS, ONE EVAL HEMODIALYSIS, REPEATED EVAL. HOLTER

I & D /DEBRIDEMENT ICD IMPLANTATION ICD INTERROGATION

ILEOSCOPY THROUGH STOMA ILEOSCOPY W/ BIOPSY INFUSION 1-8 HRS.

INFUSION TO 1 HR.

INJ FOR ANGIOGRAPHY

INJ FOR AV BYPASS GRAFTS INJ TENDON/LIGAMENT/CYST INJECTION, CARDIAC CATH INTRA-ATRIAL PACING INTRA-ATRIAL RECORDING INTRAVENTRICULAR PACING

INTRODUCTION OF NEEDLE/CATH IV FLUID THERAPY

IV INFUSION IV PUSH

IV THER. 1-8 HRS.

IV THER. UP TO 1 HR.

LASER SURGERY (NOT YAG) LEFT HEART CATHETERIZATIION

LEFT VENTRICULAR RECORDING LIVER BIOPSY

LUNG COMPLIANCE MECHANICAL VENTILATION METHACHOLINE CHALLENGE MONITOR W/ REVIEW & REPORT OVER GUIDE WIRE

PACEMAKE IMPLANTATION PACEMAKER

PACEMAKER FOLLOW UP PACEMAKER, RHYTHM STRIP PARACENTESIS

PERIPH BLOOD SMEAR INTERPRET PHLEBOTOMY

PLACE CATHETER IN VEIN, HEMO PLEURODESIS

PNEU BALLOON (30MM+) ACHALASIA PROC ABLATION (OTHER)

PROC BIOPSY

PROC CONTROL BLEEDING PROC DIAGNOSTIC (BRUSHINGS) PROC DILATION

PROC HOT BIOPSY(IES) PROC REMOVAL FB PROC SNARE

PROC TUMORS, MULT (HOT/SN/BI) PROCTOSCOPY

PROGRAMMED STIMULATION/PACING PSEUDOFOLLICULAR SCAN PULMONARY ARTERY CATHETER PULMONARY FUNCTION INTERPRET PULMONARY PROCEDURES

PULSE OXIMETRY MULTIPLE REHAB PULSE OXIMETRY SINGLE REHAB PULSE OXIMETRY, MULTIPLE RHEUMATOLOGY PROCEDURES RIGHT HEART CATHETERIZATION RIGHT VENTRICULAR RECORDING RT & LT HEART CATHETERS

SB ENDO W/ABLATION

SB ENDO W/BLEEDING CONTROL SB ENDO W/FB REMOVAL

SB ENDO W/HOT BIOPSIES SB ENDO W/INCL ILEUM

SB ENDO W/INCL ILEUM,BIOPSY SB ENDO W/INCL ILEUM,BLD CONT

SB ENDO W/TUMORS (SNARE) SCREENING, MAMMOGRAM SCREENS AND INJ, ANTI-COAG SLOW VITAL CAPACITY SMALL BOWEL ENDOSCOPY

SMALL BOWEL ENDOSCOPY,BIOPSY SOUND/BOUGIE;SINGLE/MULT SPIROMETRY

SPIROMETRY, PRE & POST STO ABLATION

STO BIOPSY

STO CONTROL HEM. STO DIAG/BRUSHING STO FOREIGN BODY STO HOT BIOPSY(IES) STO SNARE

STOMA

STRESS TEST, ECHO IMAGING STRESS TEST, EXER (NON-IMAGE) STRESS TEST, NUCLEAR IMAGING SUBCUT./IM

SYMPTOM LIMITED EXERCISE TEST THORACENTESIS

THORACIC GAS VOLUME THORACOSTOMY THRESHOLD TEST (DUAL) THRESHOLD TEST (SGL)

TILT TABLE TEST FOR SYNCOPE TRANS. BLOOD

TRANS. INDWELL. VEN. ACC. CARE TRANS. THERAPEUTIC APHERESIS TRANSFUSION

VENIPUNCTURE (ROUTINE), HEMO

# Appendix C - Instrument Processing Routines

The following is a listing of the processing routines associated with each instrument.

|  |  |
| --- | --- |
| **Instrument Name:** | **Processing Routine:** |
| CLINIVISION | MDHL7R1 |
| [30](#_bookmark73)BRAUN | MDHL7D |
| BRAUN (Bi-Directional) | MDHL7D |
| FRESENIUS | MDHL7D |
| FRESENIUS (Bi-Directional) | MDHL7D |
| GAMBRO\_EXALIS | MDHL7D |
| GAMBRO\_EXALIS (Bi-  Directional) | MDHL7D |
| Muse | MDHL7M1 |
| Muse EKG | MDHL7M1 |
| Muse Exercise | MDHL7M1 |
| Muse Holter | MDHL7M1 |
| Muse Pacemaker EKG | MDHL7M1 |
| OLYMPUS | MDHL7E |
| OLYMPUS Bronchoscopy | MDHL7E |
| OLYMPUS Colonoscopy | MDHL7E |
| OLYMPUS EGD | MDHL7E |
| OLYMPUS EGDPEG | MDHL7E |
| OLYMPUS ERCP | MDHL7E |
| OLYMPUS Endo Ultrasound | MDHL7E |
| OLYMPUS Enteroscopy | MDHL7E |
| OLYMPUS Liver Biopsy | MDHL7E |
| OLYMPUS Paracentesis | MDHL7E |
| OLYMPUS Sigmoidoscopy | MDHL7E |
| SMC | MDHL7P1 |

30 Patch MD\*1.0\*6 May 2008 Added Hemodialysis instrument entries: BRAUN, FRESENIUS, and GAMBRO.

# [31](#_bookmark75)Appendix D – Exported Values For Hemodialysis Options

## Custom Data List

### Anticoagulants

Item Value

1. Heparin
2. Citrate
3. Saline Flush
4. None
5. Warfarin

### Code Statuses

Item Value

* 1. DNR
  2. AD Signed
  3. Full Resuscitation
  4. DNI

### Dialyzer List

Item Value

01 400-HG

02 500-HG

03 50H

04 50M

05 50U

06 600-HE

07 65H

08 65U

09 700-HE

10 75U

11 90U

1. Alwall GFE-09
2. Alwall GFE-11
3. Alwall GFE-12
4. Alwall GFE-15
5. Alwall GFE-18

31 Patch MD\*1.0\*6 May 2008 Listed the exported values for Hemodialysis Options.

1. Alwall GFS Plus 11
2. Alwall GFS Plus 12
3. Alwall GFS Plus 16
4. Alwall GFS Plus 20
5. Alwall GFS-12
6. Alwall GFS-16

23 B3-0.8-A

24 B3-1.0-A

25 B3-1.0-A

26 B3-1.6-A

27 B3-2.0-A

28 BK-1.6-U

29 BK-2.1-U

30 C-061

31 C-081

32 C-101

33 C-121

34 C-151

35 CA-110

36 CA-150

37 CA-170

38 CA-210

39 CA-50

40 CA-70

41 CA-90

1. CAHP/DICEA 110G
2. CAHP/DICEA 150G
3. CAHP/DICEA 210G
4. CAHP/DICEA 90G 46 CF-12 (ST-12)

47 CF-15 (ST-15)

48 CF-23 (ST-23)

49 CF-25 (ST-25)

50 CT-110G

51 CT-190G

52 F5

53 F-50

54 F6

55 F-60

56 F-60-M

57 F8

58 F-80

59 F-80-M

1. Filtral 20
2. Lundia Alpha 400
3. Lundia Alpha 500
4. Lundia Alpha 600
5. Lundia Alpha 700
6. Lundia Aria 550
7. Lundia Aria 700
8. Lundia Pro 500
9. Lundia Pro 600
10. Lundia Pro 800

70 M-081

71 M-101

72 M-121

73 M-151

1. Optiflux 200r
2. Polyflux 11S
3. Polyflux 14S
4. Polyflux 17S
5. Polyflux 21S
6. Polyflux 210H
7. PSN120

81 PSN-150

82 PSN-170

83 PSN-210

84 T-150

85 T-175

86 T-220

1. Tricea 110G
2. Tricea 150G
3. Tricea 190G
4. Tricea 210G

### Education Codes

Item Value

01 One

### ESRD Diagnosis

Item Value

01 585

02 403.01

03 403.11

04 403.91

05 25000 A Type II, adult-onset type or unspecified type diabetes 06 25001 A Type I, juvenile type, ketosis prone diabetes

1. 5829 A Glomerulonephritis (GN)(histologically not examined)
2. 5821 A Focal glomerulosclerosis, focal sclerosing GN 09 5831 A Membranous nephropathy
3. 5832 A Membranoproliferative GN type 1. diffuse MPGN
4. 5832 C Dense deposit disease, MPGN type 2
5. 58381 B IgA mephropathy, Berger's Disease (proven by immunofluorescence)
6. 58381 C IgM nephropathy (proven by immunofluorescence)
7. 5804 B Rapidly progressive GN
8. 5834 C Goodpasture's Syndrome
9. 5800 C Post infectious GN, SBE
10. 5820 A Other proliferative GN
11. 7100 E Lupus erythematosus, (SLE nephritis)
12. 2870 A Henoch-Schonlein syndrome
13. 7101 B Sclerodema
14. 2831 A Hemolytic uremic syndrome
15. 4460 C Polateritis
16. 4464 B Wegener's granulomatosis
17. 5839 C Nephropathy due to heroin abuse and related drugs
18. 4462 A Vasculitis and its derivatives
19. 5839 B Secondary GN, other
20. 9659 A Analgesic abuse
21. 5830 B Radiation nephritis
22. 9849 A Lead nephropathy
23. 5909 A Nephropathy caused by other agents
24. 27410 A Gouty nephropathy
25. 5920 C Nephrolithlasis
26. 5996 A Acquired obstructive uropathy
27. 5900 A Chronic pyelonephritis, reflux nephropathy
28. 58389 B Chronic interstitial nephritis
29. 58089 A Acute interstitial nephritis
30. 5929 B Urolithiasis
31. 2754 A Nephrocalcinosis
32. 4039 D Renal disease due to hypertension (no primary renal disease)
33. 4401 A Renal artery stenosis
34. 59381 B Renal artery occlusion
35. 59381 E Cholesterol emboli, renal emboli
36. 75313 A Polycystic kidneys, adult type (dominant)
37. 75314 A Polycystic, infantile (recessive) 7
38. 75316 A Medullary cystic disease, including nephronophthisis
39. 7595 A Tubular sclerosis
40. 7598 A Hereditary nephritis, Alport's syndrome
41. 2700 A Cystinosis
42. 2718 B Primary oxalosis
43. 2727 A Fabry's disease
44. 7533 A Congenital nephrotic syndrome
45. 5839 D Drash syndrome, mesangial sclerosis
46. 7532 A Congenital obstructive uropathy
47. 7530 B Renal hypoplasia, dysplasia, oligonephronia
48. 7567 A Prune belly syndrome
49. 7598 B Hereditary/familial nephropathy
50. 1890 B Renal tumor (malignant)
51. 1899 A Urinary tract tumor (malignant)
52. 2230 A Renal tumor (benign)
53. 2239 A Urinary tract tumor (benign)
54. 2395 A Renal tumor (unspecified)
55. 2395 B Urinary tract tumor (unspecified)
56. 20280 A Lymphoma of kidneys
57. 2030 A multiple myeloma
58. 2030 B Light chain nephropathy
59. 2773 A Amyloidosis
60. 99680 A Complication post bone marrow or other transplant
61. 28260 A Sickle cell disease/anemia
62. 28269 A Sickle cell trait and other sickle cell (HbS/Hb other)
63. 64620 A Postpartum renal failure
64. 0429 A AIDS nephropathy
65. 8660 A Traumatic or surgical loss of kidney(s)
66. 5724 A Hepatorenal syndrome
67. 5836 A Tubular necrosis (no recovery)
68. 59389 A Other renal disorders
69. 7999 A Etiology uncertain

### Medication Routes

Item Value

* 1. ID
  2. IN
  3. IV
  4. IVP
  5. PO
  6. SL
  7. SQ

### Medication Units

Item Value

1. ml
2. mg
3. units
4. mcg
5. oz
6. gal
7. gr
8. Gm
9. Kg
10. lb
11. pt
12. in
13. qt
14. liter
15. Tsp
16. Tbsp
17. mEq

### Modalities

Item Value

* 1. HD
  2. Inpatient HD
  3. Short Intermittent HD 04 Nocturnal HD

1. ICU HD
2. Outpatient HD
3. Home HD

### TIU Note Titles

Item Value

01 Site Specific TIU Note Title

### Transportation Methods

Item Value

1. ambulatory
2. bed
3. motorized w/c
4. wheel chair
5. stretcher

## Preferences

### System Preferences

The system preferences are exported with the following default values:

Allow USER control Study Status = FALSE Allow USER delete blank F/S records = FALSE Allow USER Reset Study Status = FALSE Application Web Page URL =

<http://vista.med.va.gov/clinicalspecialties/clinproc/showProject.asp?pid=1> Blanks Placeholder = <blank>

Broker Timeout (sec) = 30 Color Disabled = -16777201 Color Editable = -16777211

Color of Background = -16777201 Color of Toolbars = 12632256 Color Read Only = 15793151 Color Read/Write = 12632256 Color Required = -16777192 Color Review = 12632256

Color Unknown = 255

Falls Assessment as Separate TIU Note = TRUE Flowsheet Refresh Rate (min) = 15

Ignore Unfinished Status = TRUE Overwrite Manual Input = TRUE

Pain assessment based on how patient tolerates pain = FALSE Pain Level = 1

Report keyword = TREATMENT REPORT Reverse Flowsheet Order = TRUE

Save Flowsheet Vitals = FALSE Save Vitals = FALSE

Set the new study Cover to Read Only = FALSE Show Additional Reports = TRUE

Show Disabled Studies to Users = FALSE Show Flowsheet Event Copies = TRUE

Show Infectious Diseases information as Tree = TRUE Show report signature field = TRUE

Show TIU Note Templates = FALSE Show Treatment Status Report = TRUE Study List Refresh Rate (sec) = 60 Study Load Limit = 5

Summary Report Name = Summary Report

## Report List

### Summary Report Template

TREATMENT REPORT for HEMODIALYSIS STUDY #<**StudyID**>

Patient Name: <**PatientName**> SSN: <**PatientSSN**>

DOB: <**PatientDOB**>

Age: <**PatientAge**>

Sex: <**PatientSex**>

Treatment Date: <**cdsInfo.StudyDate**>

ESRD Diagnosis: <**Diagnosis**>

Diagnosis Date: <**DDate**>

Initial Therapy Date:..<**InitialTDate**> Modality: <**Modality**>

Code Status: <**CodeStatus**>

Attending Nephrologist:<**Attending Nephrologi**> Schedule: <**SCHEDULE**>

Transplant Candidate. <**cdsSummary.TransCand**>

Work in Progress. <**cdsSummary.TransWIP**>

Referred to TC. <**cdsSummary.TransReff**>

Station#: <**Station**> Machine#: <**Machine**>

TREATMENT SUMMARY

Treatment Start Time:..<**Treatment Start Time**> Treatment End Time:....<**Treatment End Time**>

Treatment Duration:....<**Treatment Duration**> (instrument data) Duration Adjusted:....<**Treatment Duration M**> (manual input)

Total UF:..............<**Summary Total UF**> Total LP: <**Summary**

**Total LP**>

Mean UFR:..............<**Summary Mean UFR**> Mean TMP: <**Summary**

**Mean TMP**>

Average BFR:...........<**Summary Avg BFR**> Average DFR: <**Summary**

**Avg DFR**>

Mean Dialysis Temp:....<**Summary Mean Temp**> Mean Conductivity:.<**Summary Mean Cond**>

Total KT:..............<**Summary Total KT**> Total KT/V: <**Summary**

**Total KT/V**> URR:...................<**Summary URR**>

Intra Access BF: <**IABF**>

VP at Zero BF: <**VP0**>

AVP at Zero BF: <**AVP0BF**>

VP at 200 ml/min: <**VP200**>

Overall Comments:

<**Summary Comments**>

RX

ORDER

Dialyzer: <**cdsSummary.Dialyzer**>

Reuse: Max#: <**cdsSummary.ReuseNum**>

Tx Length: <**TxLength**>

Ultrafiltration:.<**cdsSummary.RxUltra**> kg/hr EDW:.............<**cdsSummary.RxEDW**> Kg

BFR:.............<**cdsSummary.BFR**> cc/min Dialysate Flow:..<**cdsSummary.DFlow**> cc/min Temperature:.....<**cdsSummary.Temp**> C

DIALYSATE FORMULA

K:...............<**cdsSummary.DFK**> meq/Liter

HC03:............<**cdsSummary.DFHCO3**> meq/Liter

NA:..............<**cdsSummary.DFNA**> meq/Liter

CA:..............<**cdsSummary.DFCA**> meq/Liter

ANTICOAGULANTS

Type: <**cdsSummary.ACType**>

Bolus: <**cdsSummary.ACLoad**>

Maintenance: <**cdsSummary.ACDoses**>

Duration: <**cdsSummary.ACEndTime**>

Other: <**cdsSummary.ACOther**>

MODELING NA: <**cdsSummary.MODNA**>

UF: <**cdsSummary.MODUF**>

OTHER ORDERS

<**cdsSummary.RxOther**>

Rx and Lab Notes:

<**cdsSummary.LabNotes**>

PRE-TREATMENT

WEIGHT

Pre-Weight:......<**Summary Pre Weight**> Kg Dry Weight:......<**cdsSummary.RxEDW**> Kg Goal Weight:.....<**Summary Goal Weight**> Kg

TEMPERATURE

Pre-Temp:........<**Summary Pre Temp**> F PRE-BLOOD PRESSURE AND PULSE SEATED

BP:..............<**Sum Pre BP Sys Sit**> / <**Sum Pre BP Dia Sit**> mm Hg Pulse:...........<**Sum Pre Pulse Sit**> bpm

PRE-BLOOD PRESSURE AND PULSE STANDING

BP:..............<**Sum Pre BP Sys Stand**> / <**Sum Pre BP Dia Stand**> mm Hg Pulse:...........<**Sum Pre Pulse Stand**> bpm

<**Pre Pain Report**> MENTAL STATUS

Alert:...........<**Sum Pre Alert**> Confused:........<**Sum Pre Confused**> Sedate:..........<**Sum Pre Sedate**> Unresponsive:....<**Sum Pre Unresponsive**> Lethargic:.......<**Sum Pre Lethargic**> Restless:........<**Sum Pre Restless**> Oriented:........<**Sum Pre Oriented**>

(<**Sum Pre Oriented Tex**>)

OTHER

Edema:...........<**Sum Pre Edema**> Respirations:....<**Sum Pre Resp**> Shortness of Breath: <**Sum Pre SOB**>

PATIENT EDUCATION

Has the patient been educated?...<**Educated**> Education Key:.<**EduKey**>

Education Init.<**EduInit**>

PATIENT TRANSPORTATION

Transported by:..<**PreTransportation**>

SAFETY CHECKS

Have the safety checks been performed? <**SafetyChecks**>

PRE-TREATMENT NOTES:

<**cdsSummary.PreNotes**>

ACCESS USED

<**ACCESS USED**> FLOWSHEET

<**FLOWSHEET**>

Flowsheet Notes:

<**Flowsheet Notes**>

MEDICINE ADMINISTRATION

<**MEDICINE TABLE**> POST-TREATMENT

WEIGHT

Post-Weight:.....<**Summary Post Weight**> Kg Tx Goal Weight:..<**cdsSummary.RxEDW**> Kg

TEMPERATURE

Post-Temp:.......<**Summary Post Temp**> F POST-BLOOD PRESSURE AND PULSE SEATED

BP:..............<**Sum Post BP Sys Sit**> / <**Sum Post BP Dia Sit**> mm Hg Pulse:...........<**Sum Post Pulse Sit**> bpm

POST-BLOOD PRESSURE AND PULSE STANDING

BP:..............<**Sum Post BP Sys Stan**> / <**Sum Post BP Dia Stan**> mm Hg Pulse:...........<**Sum Post Pulse Stand**> bpm

<**Post Pain Report**> MENTAL STATUS

Alert: <**cdsSummary.PostAlert**>

Confused: <**cdsSummary.PostConfu**>

Sedate: <**cdsSummary.PostSedat**>

Unresponsive: <**cdsSummary.PostUnres**>

Lethargic: <**cdsSummary.PostLetha**>

Restless: <**cdsSummary.PostRestl**>

Oriented: <**PostOriented**>

(type): <**PostOrientedText**>

OTHER

Edema: <**cdsSummary.PostEdema**>

Respirations: <**cdsSummary.PostResp**>

Shortness of Breath: <**cdsSummary.PostSOB**>

OBSERVATIONS

Was the treatment weight acheived? <**cdsSummary.POWeight**> Was any medication administered? <**cdsSummary.POMedicat**> How did the patient tolerate treatment?

Vomiting: <**cdsSummary.TlrVom**>

Hypotension: <**cdsSummary.TlrHyp**>

Syncope: <**cdsSummary.TlrSyn**>

Cramping: <**cdsSummary.TlrCram**>

Stable:........<**Sum Post Stable**> Other:.........<**cdsSummary.TlrOther**> (<**cdsSummary.TlrOtherD**>)

TRANSPORTATION

Transported by:..<**PostTransportation**>

POST-TREATMENT NOTES

<**cdsSummary.PostNotes**>

FALLS RISK EVALUATION

<**FallsAssessment**>

Report was generated by <**Version**>

at <**Now**>

# Index

#### A

alerts

setting up, 9-1 analyzer

for instruments, 6-11

Anticoagulants, 18-1

Ask Encounter Update, 9-4

#### B

background processor and test accounts, 3-6

configuring for test accounts, 3-7

Bi-Directional Interface Process Flow, 1-3 Broadcast Messages, 9-5

business rules adding, 5-1

understanding, 5-1

#### C

classes

creating, 4-3

Clinical Procedures, 1-1

flowcharts, 1-1

general information, 1-10

introduction, 1-1

populating definition file, 6-1 setting up, 6-1

troubleshooting, 13-1

Code Statuses, 18-1

Command Line Switches, 15-1 Definition, 15-1

List of, 15-3 configuring

HL7 messages, 11-1 Consult procedures

creating, 8-5 Consult Services

setting up, 8-1 Consults

setting up, 8-1

CP Gateway, 1-8

working with, 10-1

CP Manager, 1-8

CP User, 1-8

CPRS, 9-8

editing parameters. setting up, 9-1

CPRS Tools Menu, 15-2 Custom Data List, 18-1

**D**

definition file populating, 6-1

deleting an automated instrument or procedure, 2-3 Dialyzer List, 18-1

document parameters defining, 4-9

#### E

editing

instruments, 6-3

Education Codes, 18-3

ESRD Diagnosis, 18-3 Exported Procedures List, 16-1

Exported Values For Hemodialysis Options, 18-1

#### F

finding a parameter, 2-2 flowcharts

Clinical Procedures, 1-1 Force PCE Entry, 9-6

#### G

Glossary, 14-1

#### H

HL7

configuring, 11-1

fixing technical issues, 11-10 HL7 parameter

setting up, 11-1

hospital location file requirement, 1-12

#### I

images

and test accounts, 3-1 imaging, 3-1

Instrument Processing Routines, 17-1 instruments

adding, 6-8

deleting, 2-3

editing, 6-3

printing reports, 2-4

setting up, 6-2

using the instrument analyzer, 6-11 interpreter role

understanding, 5-1 introduction

Clinical Procedures, 1-1

#### M

manuals

related, 1-9

Medication Routes, 18-5

Medication Units, 18-5

Modalities, 18-6

#### N

notifications setting up, 9-1

#### P

parameters

finding, 2-2 populating

CP definition file, 6-1 Preferences, 18-7 printing

instrument reports, 2-4

procedure reports, 2-4 system parameter reports, 2-4

procedures

adding, 6-16

deleting, 2-3

editing, 6-12

printing reports, 2-4

setting up, 6-12

#### R

related manuals, 1-9

Report List, 18-9 reports

printing, 2-4

resource requirements, 1-11

#### S

share folder

configuration for an automated instrument, 12-1

Shared Broker Environment, 15-1 Summary Report Template, 18-9 system parameter reports

printing, 2-4 system parameters

allow non-instrument attachments, 6-19 bypass CRC checking, 6-20

calculating a file’s CRC value, 6-22 Clinical Procedures home page, 6-20 Clinical Procedures on-line, 6-20 CP/BGP Transfer Directory, 6-20 CRC values, 6-21

days to keep instrument data, 6-22 imaging file types, 6-23

offline message, 6-24

setting up, 6-18

version compatibility, 6-24

V*IST*A scratch HFS directory, 6-26 System Preferences, 18-7

#### T

test accounts, 3-1

changing, 3-2

configuring background processor, 3-7 refreshing, 3-10

working with background processor, 3-6 titles

creating, 4-3 TIU

setting up, 4-1 toolbar, CP Manager, 2-1

Transportation Methods, 18-6

troubleshooting, 13-1

#### U

Uni-Directional Interface Process Flow, 1-5

#### W

workload

implementing, 1-12

reporting, 1-12