LABORATORY HEWLET PACKARD (HP)

CAREVUE INTERFACE

Patch LR*5.2*254 and Patch LA*5.2*53

INTERFACE SPECIFICATIONS

Version 5.2

May 2000

Department of Veteran Affairs

VISTA Software Service

Clinical Specialties Product Line

# Introduction

The purpose of this document is to describe the Veterans Health Information Systems and Technology Architecture (VISTA) Laboratory Test result system’s use of the HL7 Interface Standard (Version 2.3) and to provide a reference guide for implementing this interface with the CareVue system.

The focus of this document is to describe the HL7 Observation Request Unsolicited (ORU) control message used for sending VISTA lab test results to the CareVue system. The interface is designed to generate an ORU message upon verification of lab results for any patient whose reporting location is identified as either a CareVue location.

# Overview - VISTA Laboratory Interface

The VISTA Laboratory interface provides the capability for VISTA to broadcast laboratory results that are entered into VISTA. Each time laboratory results are verified in VISTA, the laboratory interface determines if those results should be sent to the Interface Engine (see paragraph below).

Figure 1: VISTA Laboratory Interface

## VISTA Filtering of Results

Not all laboratory results are sent to the Interface Engine. Each time laboratory results are verified in VISTA, the location designated as the report routing location (this defaults to the ordering location entered at the time of ordering the laboratory tests in VISTA) is checked to see if the results should be sent to the CareVue. VISTA uses the FOREIGN INTERFACE file (#62.487) for this purpose.

If there is an entry in this file for the specified location, an HL7 message is generated and sent to the Interface Engine. The location must match exactly an entry in the FOREIGN INTERFACE and must be a valid location in the Hospital Location file (#44).

# VISTA System Environment

VISTA uses standardized HL7 messaging software, HLLP, for communicating with external systems. Within VISTA, different instances of the HLLP software (VMS processes) are associated with different applications. Each instance of the HLLP software is assigned a unique name in the VISTA device file.

Table 1 identifies the HLLP process names associated with the HOST integration.

## HLLP Operation - VISTA Results Interface

VISTA provides the optional capability to automatically restart HLLP’s each time the running process is terminated (because of an interruption on the communications connection with the Interface Engine or other system interruption). This option should be selected for each HLLP communicating with the Interface Engine.

Table 1: VISTA HLLP Processes - HOST Integration Example

| VISTA Interface | VISTA HLLP Process | VISTA Virtual Device |
| --- | --- | --- |
| VISTA Outgoing Interfaces (VISTA to Interface Engine) |
| ADT - Test | OG_TEST | LTA_9297 |
| ADT - Production | OG_PROD | LTA_9299 |
| Laboratory Results - Test | OG_LAB_TEST | LTA_9333 |
| Laboratory Results - Production | OG_LAB_PROD | LTA_9334 |
| VISTA Incoming Interfaces (Interface Engine to VISTA) |
| Results - Test | IC_TEST | LTA_9161 |
| Results - Production | IC_PROD | LTA_9296 |

# VISTA Network Connectivity to Interface Engine

The VISTA HL7 software uses asynchronous type communications. Separate communications lines have been dedicated for each of the VISTA interfaces that have been implemented with the Interface Engine. Configuring the system to support both a test interface and a production interface requires an additional set of dedicated lines.

The communication lines between VISTA and the Interface Engine are implemented by dedicating ports on a Digital terminal server. Each of the lines requires two terminal server ports, one for the VISTA end of the communication and one for the Interface Engine end of the communication. The two terminal server ports are physically connected using a null modem cable.

Both VISTA and the Interface Engine communicate with the terminal server(s) using the Local Area Transport (LAT) protocol over the Ethernet backbone.

Figure 2 illustrates how the connections between VISTA and the Interface Engine are implemented.

Figure 2: VISTA to Interface Engine Network Connections

These ports can be provided by any terminal server(s) connected to the network with the only requirement being that each pair (VISTA end and Interface Engine end) be physically connected using a null modem cable.

Table 2 identifies the terminal server ports, which are being used to communicate between VISTA and the Interface Engine.

Table 2: Terminal Server Ports - VISTA to Interface Engine Communications Example

| VISTA Interface HL7 LLP | Terminal Server Port (VISTA Connection) VISTA Virtual Device | Terminal Server Port (Interface Engine Connection) Interface Engine Virtual Device |
| --- | --- | --- |
| Laboratory Results I/F - Production INTERFACE ENGINELABRESULTS LLP | DSV 104 Port 2 LTA_9334 OG_LAB_PROD | DSV 104 Port 1 LAT device 649 vista_ic_lab |

Laboratory Results Interface (vista_ic_lab)

The laboratory interface is an outgoing interface from VISTA to the Interface Engine which transmits HL7 ORU messages to communicate laboratory results that have been verified in VISTA. The Interface Engine translates these incoming messages for communication to HP CareVue.

Communications lines and terminal server port designations are assigned in VISTA by the IRM operations staff and are defined as follows:

Outgoing

NAME: OG_LAB_PROD                      $I: _LTA9334:

LOCATION OF TERMINAL: COMPUTER ROOM   ASK HOST FILE: YES

SUBTYPE: C-OTHER                      TYPE: TERMINAL

LAT SERVER NODE: DSV104               LAT SERVER PORT: PORT_6

Incoming

NAME: IC_LAB_PROD                      $I: _LTA9296:

LOCATION OF TERMINAL: IRM COMPUTER ROOM

SUBTYPE: P-OTHER                      TYPE: TERMINAL

LAT SERVER NODE: DSV104               LAT SERVER PORT: PORT_8

Communication between VISTA and the Interface Engine is accomplished using an asynchronous serial communications protocol as if the systems were connected using a standard RS232 asynchronous serial communications line.

Table 3 and Table 4 identify the terminal server port settings for the VISTA interface with the Interface Engine. These settings are used for terminal server ports, one for the VISTA end, and the other for the Interface Engine end. IRM operations staff can modify terminal server port settings.

Table 3: Terminal Server Port Settings - VISTA to Interface Engine ADT Interface

| Port n |  | Server: DSVnnn |
| --- | --- | --- |
| Character Size: | 8 | Input Speed: | 9600 |
| Flow Control: | XON | Output Speed: | 9600 |
| Parity: | None | Signal Control | Enabled |
| Stop Bits: | Dynamic | Signal Select: | CTS-DSR-RTS-DTR |
| Access: | Remote | Local Switch: | None |
| Backwards Switch: | None | Name: | PORT_n |
| Break: | Remote | Session Limit | 1 |
| Forwards Switch: | None | Type: | Soft |
| Default Protocol: | LAT |  |  |
| Preferred Service: None |  |  |
| Authorized Groups: 0 (Current) Groups: 0 |  |  |
| Enabled Characteristics: Autoconnect, DTRwait |  |  |

Table 4: Terminal Server Port Settings - VISTA to Interface Engine Results Interface

| Port n |  | Server: DSVnnn |
| --- | --- | --- |
| Character Size: | 8 | Input Speed: | 9600 |
| Flow Control: | XON | Output Speed: | 9600 |
| Parity: | None | Signal Control | Enabled |
| Stop Bits: | Dynamic | Signal Select: | CTS-DSR-RTS-DTR |
| Access: | Remote | Local Switch: | None |
| Backwards Switch: | None | Name: | PORT_n |
| Break: | Remote | Session Limit | 1 |
| Forwards Switch: | None | Type: | Soft |
| Default Protocol: | LAT |  |  |
| Preferred Service: None |  |  |
| Authorized Groups: 0 (Current) Groups: 0 |  |  |
| Enabled Characteristics: Autoconnect, DTRwait |  |  |

# Detailed Description - VISTA Laboratory Results Interface

VISTA sends laboratory results messages to the INTERFACE ENGINE using an HL7 message format. The Interface Engine converts these messages to the message formats required by the CareVue systems and sends the messages to CareVue.

The VISTA Laboratory Results Interface Mapping describes the messages according to the HL7 format that is received from VISTA.

# VISTA Laboratory Results Interface Mapping

Laboratory results are sent to the Interface Engine as HL7 ORU Unsolicited Results messages. Table 5 identifies the HL7 message segments that comprise the message format for result messages being sent from VISTA.

Table 5: ORU Message Layout

| HL7 Segment | Description |
| --- | --- |
| MSH | Message Header Segment |
| PID | Patient Identification Segment |
| OBR | Observation Request Segment |
| NTE | Optionally, one or more Notes and Comments Segments. |
| OBX | One or more Observation Segments |

The following tables identify the specific data elements for each of the HL7 segments. These tables identify only those fields that are supported by the VISTA interface. All other fields in the HL7 segments are not currently used by VISTA.

Note: VISTA uses HL7 control characters in a different manner than defined in the current HL7 Standard.

Table 6: MSH- Message Header Segment

| HL7 Seq | HL7 Field Name | Req’d or Opt | Data Type | Length | VISTA Value/ Comments |
| --- | --- | --- | --- | --- | --- |
| 1 | Field Separator | R | ST | 1 | “^” see notes below |
| 2 | Encoding Characters | R | ST | 4 | “~¦\&” see notes below |
| 3 | Sending Application |  | ST | 15 | VISTA sends as “DHCP CARELIFE” |
| 4 | Sending Facility |  | ST | 20 | VISTA sends station number |
| 5 | Receiving Application |  | ST | 30 | VISTA sends as “INTERFACE ENGINE CARELIFE” |
| 6 | Receiving Facility |  | ST | 30 | VISTA sends station number |
| 7 | Date/Time of Message |  | TS | 26 | YYYYMMDDHHMMSS |
| 8 | Security |  | NM | 40 | Not used |
| 9 | Message Type | R | CM | 7 | VISTA sends as “ORU^R01” see notes below |
| 10 | Message Control ID | R | ST | 20 | Unique message ID generated by the sending system |
| 11 | Processing ID | R | ID | 1 | VISTA sends as “P” see notes below |
| 12 | Version ID | R | ID | 8 | VISTA sends as “2.3” |
| 15 | Accept Acknowledgment Type |  | ID | 2 | VISTA sends as “AL” (Always) |
| 16 | Application Acknowledgment Type |  | ID | 2 | VISTA sends as “AL” (Always) |
| 17 | Country Code |  | ID | 2 | VISTA sends as “USA” |

MSH Segment Notes

Data Type: ST: String data - left justified with optional trailing blanks
TS: time stamp always in the format YYYYMMDDHHMM[SS][+/-ZZZZ]
CM: A field that is the combination of other meaningful data fields.
ID: ST formatted data that corresponds to a table entry maintained by system

NM: Numeric

Field Separator & Encoding Characters:

| VISTA Value | Description |
| --- | --- |
| ^ | field separator |
| ~ | component separator |
| ¦ | repetition separator |
| \ | escape character |
| & | subcomponent separator |

Message Type:  <message type><trigger event>

Processing ID:  “D” - Debugging messages sent from INTERFACE ENGINE Test Environment
“P” - Production messages sent from INTERFACE ENGINE Production Environment

Table 7: PID – Patient Identification Segment

| HL7 Seq | HL7 Field Name | Req’d or Opt | Data Type | Length | VISTA Value / Comment |
| --- | --- | --- | --- | --- | --- |
| 3 | Patient ID (Internal ID) | R | CM | 20 | VISTA internal entry number. |
| 6 | Mother’s Maiden Name |  | ST | 48 | VISTA Mother’s maiden Name |

PID Segment Table Notes:

Data TypeCM: A field that is the combination of other meaningful data fields.
 ST: String data - left justified with optional trailing blanks

Patient ID (Internal ID): <patient ID (ST)><check digit (NM)><check digit scheme (ID)>
                            <assigning facility ID (ST)><type (ID)>

Note:  VISTA calculates the check digit as the check digit scheme applied to the sum of the digits, where non-numeric digits are calculated as zero (0).
This check digit scheme DOES NOT conform to the HL7 standard.

Table 8: OBR – Observation Request Segment

| HL7 Seq | HL7 Field Name | Req’d or Opt | Data Type | Length | VISTA Value / Comment |
| --- | --- | --- | --- | --- | --- |
| 3 | Filler Order Number | R | CM | 75 | VISTA Accession Number see notes below |
| 4 | Universal Service ID | R | CE | 200 | VISTA copies the accession area identification from Filler Order Number into the 1st component. see note below |
| 7 | Observation Date/Time |  | TS | 26 | YYYYMMDDHHMMSS |
| 14 | Specimen Received Date/Time |  | TS | 26 | YYYYMMDDHHMMSS |
| 15 | Specimen Source |  | CM | 300 | VISTA limits to 150 characters. see notes below |
| 22 | Results Rpt/Status Chng - Date/Time |  | TS | 26 | Date/Time results are verified in VISTA. |
| 24 | Diagnostic Serv Sect ID |  | ID | 10 | VISTA sends as “LAB” |
| 28 | Result Copies To |  | CN | 150 | VISTA sends the location specified in the order as the 1st component. |

OBR Segment Table Notes:

Data Type:  CM: composite field which may include multiple components
CE:  Coded Element - contains six components in two groups as follows:
        <identifier><text><name of coding system>
        <alternate identifier><alternate text><name of alternate coding system>
TS: time stamp always in the format YYYYMMDDHHMM[SS][+/-ZZZZ]
ID: ST formatted data that corresponds to a table entry maintained by system

Filler Order Number: <unique filler ID><filler application ID>

<unique filler ID> contains VISTA accession number in the format
 xxxx nnnn aaaa, where

xxxx is a numerical prefix whose value depends on when (day (mmdd),
         month (mm) or year (yy)) the numeric accession number
         resets to 1.
nnnn is the numeric accession number and
aaaa is the VISTA accession area

<filler application ID> is not sent

Specimen Source:  <specimen source (CE)><additives (TX)><free text (TX)><body site (CE)>
<site modifier>, where VISTA sends specimen source a	s: 
<SNOMED Code><Topography Name>

Universal Service ID:

<AU~AUTOPSY~L>

<BB~BLOOD BANK~L>

<CH~CHEMISTRY~L> or any other name

<CY~CYTOPATHOLOGY~L>

<EM~ELECTRON MICROSCOPY~L>

<MI~MICROBIOLOG~L>

<SP~SURGICAL PATHOLOGY~L>

Specimen Source:   <SNOMED code1 ;Topography name1~ SNOMED code2 ;Topography name2>

Example:  <0X400;Plasma~0X400;Plasma>

Table 9: NTE – Notes and Comments Segment

| HL7 Seq | HL7 Field Name | Req’d or Opt | Data Type | Length | VISTA Value / Comment |
| --- | --- | --- | --- | --- | --- |
| 2 | Source of Comment |  | ID | 8 | VISTA sends as “L” |
| 3 | Comment |  | TX | 64K | VISTA sends one line of comment text in each NTE segment and allows multiple segments. |

NTE Segment Table Notes:

Data Type:  ID: Value corresponds to a table entry maintained by system 
TX: String data (including leading spaces)

Table 10: OBX – Result Segment - VISTA Laboratory Interface

| HL7 Seq | HL7 Field Name | Req’d or Opt | Data Type | Length | VISTA Value / Comment |
| --- | --- | --- | --- | --- | --- |
| 2 | Value Type | R | ID | 2 | VISTA does not send. |
| 3 | Observation Identifier | R | CE | 80 | VISTA sends the lab test name as the 2nd component. |
| 5 | Observation Results |  | ST | 64K | Verified results VISTA limits to 20 characters |
| 6 | Units |  | CE | 20 | Verified units used in test, e.g. mg/dL, %, NEG-POS etc. |
| 7 | References Range |  | ST | 40 | “reference low - reference high” “critical low - critical high” “therapeutic low - therapeutic high” |
| 8 | Abnormal Flags |  | ID | 10 | VISTA sends as: “_” or “N” = Normal “H” = High “HH” = Dangerously High “L” = Low “LL” = Dangerously Low |
| 11 | Observe Result Status | R | ID | 2 | VISTA only sends results with a status of “F” = Final. Corrected results are also sent with a status of “F” = Final. |

OBX Segment Table Notes:

Data Type:  ID: Value corresponds to a table entry maintained by system 
CE:  Coded Element - contains six components in two groups as follows:
        <identifier><text><name of coding system>
        <alternate identifier><alternate text><name of alternate coding system>
ST: String data
CM: Composite field which may include multiple components.

Table 11: MSA – Message Acknowledgement Segment

| HL7 Seq | HL7 Field Name | Req’d or Opt | Data Type | Length | VISTA Value / Comment |
| --- | --- | --- | --- | --- | --- |
| 1 | Acknowledgment Code | R | ID | 2 | “AA”, “AE”, “AR” |
| 2 | Message Control ID | R | ST | 20 | Message Control ID of the sending system. It is echoed back by the receiving system. |
| 3 | Text Message |  | ST | 80 | An optional text field that further describes an error condition. |

# Text boxes

Interface

Engine

HP

CareVue

Laboratory

VISTA

Laboratory