# ATTACHMENT C

Biological Safety Cabinets:  Description of and appropriate Biosafety Levels for use.

# Adapted from:“Biosafety in Biomedical and Microbiological Laboratories”; U.S. Department of Health and Human Services, National Institutes of Health, 5th Edition (2007).

# Comparison of Biological Safety Cabinets

|  |  |  | Applications |  |
| --- | --- | --- | --- | --- |
| BSC Class | Face Velocity | Airflow Pattern | Nonvolatile Toxic Chemicals and Radionuclides | Volatile Toxic Chemicals and Radionuclides |
| I | 75 | In at front through HEPA to the outside or into the room through HEPA | Yes | When exhausted outdoors 1, 2 |
| II, A1 | 75 | 70% recirculated to the cabinet work area through HEPA; 30% balance can be exhausted through HEPA back into the room or to outside through a canopy unit | Yes (minute amounts) | No |
| II, B1 | 100 | 30% recirculated, 70% exhausted. Exhaust cabinet air must pass through a dedicated duct to the outside through a HEPA filter | Yes | Yes (minute amounts) 1, 2 |
| II, B2 | 100 | No recirculation; total exhaust to the outside through a HEPA filter | Yes | Yes (small amounts) 1, 2 |
| II, A2 | 100 | Similar to II, A1, but has 100 lfpm intake air velocity and plenums are under negative pressure to room; exhaust air can be ducted to outside through a canopy unit | Yes | When exhausted outdoors (Formerly “B3”) (minute amounts) 1, 2 |
| III | N/A | Supply air is HEPA filtered. Exhaust air passes through two HEPA filters in series and is exhausted to the outside via hard connection | Yes | Yes (small amounts) 1, 2 |
| Installation may require a special duct to the outside, an in-line charcoal filter and a spark proof (explosion proof) motor and other electrical components in the cabinet. Discharge of a Class I or Class II, Type A2 cabinet into a room should not occur if volatile chemicals are used |  |  |  |  |
| In no instance should the chemical concentration approach the lower explosion limits of the compounds. |  |  |  |  |