# VBECS 2.3.3 Known Defects and Anomalies

# *Introduction*

The Known Defects and Anomalies (KDAs) table consists of system actions that do not meet performance expectations established in VBECS design documents. Some defects and anomalies require user workarounds such as being directed to view information in one report that was expected in another report. Users may refer to the table when troubleshooting issues experienced at their local facility. Having an issue known and documented on the table does not prevent users from entering a service request with the Service Desk for assistance. Entering a service request for defects and anomalies important to you helps to prioritize what gets fixed in future patch releases.

The list of KDAs is maintained for the VBECS product and updated as needed with each VBECS release as new KDAs are identified or existing KDAs are fixed.

### *Changes for VBECS 2.3.3 Rev B*

This table highlights the changes to the KDA list for the VBECS 2.3.3 Rev B release.

| **Added or Removed** | **VBECS Option** | **Reason** |
| --- | --- | --- |
| Added | VBECS Admin 86447 | Identified by a VBECS customer. |

### *Changes for VBECS 2.3.3 Rev A*

This table highlights the changes to the KDA list for the VBECS 2.3.3 release.

| **Added or Removed** | **VBECS Option** | **Reason** |
| --- | --- | --- |
| Added | CPRS 64994 | Identified by the VBECS team. |
| Added | Incoming Shipment 64944 | Identified by a VBECS customer. |
| Added | Patient Testing: Antigen Typing Test 75709 | Identified by a VBECS customer. |
| Added | Patient Testing: Record a Crossmatch 84332 | Identified by the VBECS team. |
| Added | Select Units for a Patient 70667 | Identified by a VBECS customer. |
| Added | Select Units for a Patient 84872 | Identified by the VBECS team. |

## Risk Assessment and Impact to Patient Care Assessment

All reported defects and anomalies are assessed by the VBECS team for potential harm and the impact to patient care. See Figure 1: Risk Assessment Table for the relationship between the Likelihood of Occurrence, the Level of Concern and the resulting Risk.

All defects and anomalies in this KDA must have an associated Risk rating of “Acceptable” and a Level of Concern of “Minor” indicating that there is no expectation of injury to the patient, operator, or bystander as a result of software failure, including the possible application of a mitigating workaround. The Likelihood of Occurrencecategories are Frequent, Probable, Occasional, Remote, or Improbable, as indicated by the business process.

**Important Note:** All system errors/shutdowns occur where the user is normally prohibited from proceeding to process the unit or patient in VBECS; training users to STOP and evaluate the correctness of continuing their action manually is strongly recommended.

**Figure 1: Risk Assessment Table:**

|  |  |
| --- | --- |
|  | **Levels of Concern** |
| **Minor** | **Moderate** | **Major** |
| **Likelihood of Occurrence** | **Frequent** | Acceptable | Intolerable | Intolerable |
| **Probable** | Acceptable | Intolerable | Intolerable |
| **Occasional** | Acceptable | Intolerable | Intolerable |
| **Remote** | Acceptable | ALARP | Intolerable |
| **Improbable** | Acceptable | Acceptable | ALARP |

## Related Manuals and Materials

* VBECS 2.3.3 Technical Manual-Security Guide
* VBECS 2.3.3 Administrator User Guide
* VBECS 2.3.3 User Guide
* VBECS Frequently Asked Questions (FAQS) on the VBECS SharePoint site

## How the Known Defects and Anomalies is Organized

* The table is organized by the option where the issue occurs in VBECS. “Throughout VBECS” is the only section where the item may occur in various places within the application and not only in one option.
* Description of the Issue, Recommended Workaround, if any, and Likelihood of Occurrence columns provide pertinent information about the defect or anomaly.

### *Terms*

See the VBECS User Guide Glossary for definitions of other terms and acronyms used in this table.

### *Known Defects and Anomalies*

| **VBECS Option** | **Description** | **Recommended Workaround** | **Likelihood of Occurrence** |
| --- | --- | --- | --- |
| ABO/Rh Confirmation70543 | A user entering retrospective results for ABO/Rh Confirmation always sees a "Rack QC Not Performed" override regardless of whether rack was QC'd that day or not. | Verify QC was performed, then enter a comment and proceed. | Occasional |
| Accept Orders: Pending Order List64591 | A system crash occurs if you try to filter orders by date range or specimen and enter a patient ID in a format other than Full SSN, Last-4, or Last Initial and Last-4. | Enter patient IDs using the approved formats: Full SSN, Last-4, or Last Initial and Last-4. | Remote |
| Accept Orders: Pending Order List65083 | An expired order warning message displays incorrectly based on the date the component order was received (first displays on the Accept Orders Pending Order List) by VBECS, not the date the order was processed and accepted in VBECS (accepted and moved to the Component Order Pending Task List). | Click through the message and proceed. | Occasional |
| Audit Trail Report64554 | The Audit Trail Report groups reagent level changes on the same day without regard to case-sensitivity. For example, Minimum Reagent Level changes for both K and k on the same day will be grouped under a single heading on the Audit Trail Report. | Update the Minimum Reagent Levels on different days and print the report on each day. | Occasional |
| Audit Trail Report70695 | 2mL RBC Contamination change is not documented on the Audit Trail Report. | Changes made to the >=2mL RBC Contamination field are displayed in the Incoming Shipment section of the Unit History Report. | Occasional |
| Audit Trail ReportUnit History Report64194 | When the user inactivates a unit’s ABO/Rh confirmation test results or inactivates a unit record and enters the required comment, VBECS does not print the comment on the Audit Trail Report or Unit History Report. | Print the Audit Trail Report for this activity and manually complete the comment; save it for review. The inactivation is recorded and maintained. The unsaved comment does not impact patient testing or transfusion records. | Frequent |
| Automated Testing70403 | Clicking on the Order Reflex button from the Automated Testing window for a XM will cause a system error or crash. | Order Reflex Tests using the main menu option (Orders, Order Reflex Test). | Remote |
| Automated Testing70764 | A parser error is displayed when a partial TAS message is sent to VBECS from an automated testing instrument. | Wait until both the ABO/Rh and ABS testing is complete before performing TAS on the instrument. Only send results to VBECS after both ABO/Rh and ABS results are marked as reviewed so that the message contains both tests. | Occasional |
| Cost Accounting Report63767 | When a user places a unit on the outgoing shipment invoice, cancels the invoice, and then places the same unit on another outgoing shipment invoice, VBECS displays a return credit twice on the report. | Deselect a unit prior to canceling the invoice to avoid the credit displaying twice on the Cost Accounting Report. | Occasional |
| Cost Accounting Report70841 | When shipments are entered retrospectively for a newly entered collection facility, units with an invoice date prior to the time the new collection facility record is added do not appear on the cost accounting report. | When entering shipments retrospectively, do not select a date of invoice that falls before the collection facility activation date. Always activate the collection facility first and then create a shipment with a date of shipment following that activation date. | Improbable |
| CPRS64994 | Orders that expire in VBECS are updated in CPRS but the laboratory order component is not marked as expired. | The expired order can be cancelled which causes CPRS to show it as 'discontinued' rather than 'active' or 'expired'. | Frequent |
| CPRS: Email Alert70738 | When an order is placed in CPRS, but cancelled before the order is accessioned, VBECS will send an email alert with the following text: An Exception occurred trying to process a cancel order request from CPRS.The order does cancel in CPRS, but still displays in the Pending Order List in VBECS. | The email alert can be ignored. The pending order can be canceled in VBECS, without issue. | Probable |
| Discard or Quarantine Units64863 | A VBECS system error occurs when a user tries to discard a unit that was brought in through Incoming Shipment with an inactive shipper. | Activate the shipper and then discard the unit as desired. VBECS does not allow a user to save a unit without an active shipper unless the VistALink connection is lost during unit entry. | Occasional |
| Discard or Quarantine Units70223 | VBECS will allow issued units to be discarded when the Shipment sub-menu is used and the invoice also contains other in-date non-issued units. Once an issued unit is discarded it must be corrected via a service desk ticket. | Use the Blood Units: Discard Quarantine menu option to discard units. | Occasional |
| Document ABO Incompatible Transfusion65067 | When a unit is scanned or typed, setting the volume (mL) field results in a system crash. This only occurs if the Transfusion Interrupted option is selected. | Use the search button to select the unit to be transfused. | Occasional |
| Enter Daily QC Results64889 | Reagents that are not associated with test results when the QC is partially saved are marked as satisfactory on the Testing Worklist Report. When the user partially saves rack QC results and finishes it later, the results that were originally left blank and are now completed display as unsatisfactory on the Testing Worklist Report even though the reagent displays as satisfactory in Enter Daily Reagent QC | Complete all QC testing in one instance, saving once.Do not perform partially completed QC testing. | Occasional |
| Enter Daily QC Results64986 | Changing the tested with lot number (ex. LISS) does not enable the Enter Daily Reagent QC test grid. | Change the lot number of the primary reagent as well as the secondary reagent to allow testing of both. Refer to FAQ: Retesting QC for a listing of primary and secondary reagents. | Occasional |
| Enter Daily QC Results63986 | Partial QC results cannot be saved if results are not entered for all testing phases for a given reagent. | Enter results for all testing phases for a given reagent. | Occasional |
| Enter Daily QC Results63998 | An error occurs when performing daily QC on multiple racks when at least one of the racks is already partially QC'd. | Perform QC for only one rack at a time or for multiple racks to segregate non-QC’d racks from partially completed racks. Normal workflow is to complete all QC testing together or to have each user perform their own QC. | Occasional |
| Enter Daily QC Results63724 | When a user clicks No on the decision box to not use a reagent/antiserum that is within 24 hours of expiration (23:59 on the expiration date), VBECS moves the cursor to the next reagent lot number field and does not remove the lot number of the previous cell. | Change the lot number of the reagent or continue using it until the actual expiration date and time. | Occasional |
| Enter Daily QC Results67896 | Polyspecific AHG always displays on the lot number page though it is may not be used. | Enter the Lot number of the specific AHG used for antibody screen QC. Note in the procedure that the lot number is accurate for the reagent used in QC and testing. | Occasional |
| Enter Daily QC Results65076 | When performing Daily QC with an expired reagent, the user gets no override warning for the expired reagent. No Exception report is captured. | VBECS displays that the Reagent is Expired by marking it with a red E. Do not select and use expired reagents. When expired reagent must be selected, print the Daily Reagent QC Testing Worklist report and add a manual override comment regarding the expired reagent testing details. | Occasional |
| Enter Daily QC Results6517565234 | When entering results on a partially tested QC rack, the exception "Decrease in reagent reactivity of 2 or more" does not occur.Previous QC results for Screening Cells and ABO Reverse Typing reagents are not being recognized properly to identify significant changes in reactivity (>2+). The system does not display an override associated with “Decrease in Reagent Reactivity". | The tech checks the previous day's Testing Worklist Report to view the reagent’s reactivity and assess it to avoid reagent reactivity problems. Supervisor review of daily testing and quality control reactivity is recommended within 24 hrs. Compare the results from the previous date's QC manually. | Occasional |
| Exception Report6349063500 | The “Expired reagent QC’d” exception type section contains blanks in the Rack and Phase columns. | Manually write information from the Testing Worklist Report on the printed Exception Report. | Occasional |
| Exception Report63892 | The override exception comment is not saved when invalidating Crossmatch results. | Enter any comments into the comment field of the test grid prior to invalidating Crossmatch results. | Occasional |
| Exception Report63905 | When patient ABO/Rh results are not entered in the order of performance (current testing is entered before the retrospective data entry), the Exception Report entries for an ABO/Rh interpretation discrepancy are displayed based on the time the data are saved.  | Refer to the Testing Worklist Report which displays the data in the correct chronological order. | Occasional |
| Finalize/ Print TRW64686 | When editing a TRW to change the comment for an implicated unit, only the comment detail field is saved. | Submit a service desk ticket if the comment shown on the report needs to be modified. | Occasional |
| Finalize/ Print TRW64080 | VBECS only displays 350 characters entered into the Transfusion Reaction Details field on the Finalized Transfusion Reaction Report.  | The details text is available to ~1000 characters on the Transfusion Reaction Count Report (Detailed). The printed report can be updated manually. | Occasional |
| Free Directed Unit63717 | The user displayed in the Override column for the **Release Patient Restriction** exception is the user that entered the data, regardless of which name was selected in the Removed By field. | Record the Removed By user in the comment field. | Occasional |
| Free Directed Unit64110 | A system crash occurs in Free Directed Unit when a unit conflicts with an existing record. For example, freeing W000000001 E0124D00 when there is already a W000000001 E0142V00 in inventory | When a unit is received from the blood supplier with a donation type of “V” and the unit is to be restricted to a patient as a directed donation, the product code must be manually entered with a “D” to allow the restriction in a Full Service Blood Bank. When the user tries to bring the unit into inventory using both methods, the incorrect unit can be invalidated to prevent the duplicate record. The unit label will reflect the correct unit status per local policy. A Transfusion Only facility type will not encounter this problem as the option Free Directed Unit is not enabled. | Occasional |
| Incoming Shipment64944 | VBECS does not allow the entry of a 10-digit FDA number in Incoming Shipment when the user is prompted to activate the facility. | Close the Incoming Shipment window and open Tool, Local Facilities to create a new supplier with a 10-digit FDA number. | Remote |
| Invalidate Patient Test Results65048 | A comment is required when invalidating a patient test. If the spacebar is pressed after selecting a comment, the comment is cleared and VBECS still allows the invalidation to be saved with a missing/blank comment by clicking OK.If test results are invalidated with a missing/blank comment, VBECS will crash when opening the Special Instructions & Transfusion Requirements option for the affected patient.  | Review the comment before saving to confirm the comment is correct.If a test invalidation is saved with a missing/blank comment, file a service desk ticket to have data corrections made. | Occasional |
| Invalidate Patient Test Results63787 | When the same unit is assigned, crossmatched, released, reassigned, and re-crossmatched on the same specimen, the invalidate test results window will show two (2) XM entries. If one of the XM tests invalidated, the system will crash when the remaining XM test is selected. | Do not click on the second XM test. Invalidating the first XM entry will invalidate both XM tests. | Occasional |
| Issue Blood Components70808 | A supervisor receives an error when trying to override the issue of a unit to a patient with two discrepant ABO/Rh results that have not been justified.  | Justify the patient's ABO/Rh change or invalidate the last ABO/Rh result that created the discrepancy. | Remote |
| Issue Blood Components70849 | When issuing a unit the user can set the issue date & time to fall before the unit was entered into VBECS and assigned to the patient. VBECS currently does not give the user a warning of this situation at the point of entering the date & time. | Verify that the entered date of issue falls after the blood unit was entered into VBECS and selected for the patient. | Remote |
| Log In Reagents63738 | After a reagent has been added to an invoice, that reagent is selectable from the Inventory list to allow for data corrections. If a Comment and Details text were previously entered, they are now combined into the Details field. | Reselect the desired comment and correct the Details text.Note: The Inventory list will not reflect the update. | Occasional |
| Log In Reagents64360 | VBECS crashes when the Inventory List by Invoice is sorted by clicking the Invoice column header if any of the invoices contain non-numeric characters. | Do not sort the Inventory List by Invoice. | Occasional |
| Log In Reagents63924 | Entering a decimal point in the Vials Received per Lot Number field causes unexpected behavior. | Enter whole numbers; do not enter decimals in the Vials Received per Lot Number field. Check the accuracy of the entry before saving. | Occasional |
| Log In Reagents/ Update Reagent Inventory69148 | Comments entered for unsatisfactory reagents are not viewable on any window or report. | The comments are saved in VBECS but are not viewable. When the comments must be verified or viewed, the user can file a request for this information with the Service Desk. | Frequent |
| Maintain Equipment65206 | New Maintenance Events cannot be immediately edited without first saving the record. | If changes need to be made to a newly created Maintenance Event, click OK to save it first before making the changes. | Occasional |
| Maintain Minimum Reagent Levels63550 | Entering a decimal point in the Minimum Stock Level field causes unexpected behavior. | Do not enter decimals in the Minimum Stock Level field. Check the accuracy of the entry before saving. | Occasional |
| Modify Units (not Pool or Split)70699 | When modifying a unit that has RBC Antigens, the target area does not display the RBC Antigens. This is a display issue only; the underlying data are correct. | Press the tab key several times until the target area populates correctly. | Remote |
| Order History Report64808 | In the calendar control, when the "previous month" button is clicked to set the end date earlier than the start date, VBECS will present an infinite loop of error messages. This only happens when setting the date range for an Order Summary Report found under Reports, Orders History Report. | Use the keyboard to set the date range fields.Once the endless errors start you must Sign Out (not disconnect) the VBECS server to make it stop. | Occasional |
| Order Reflex Tests68630 | VBECS does not allow a user to reflex to a weak D or patient antigen typing test from an ABO/Rh test. | Order a reflex antibody ID (ABID) from the ABO/Rh test. Then open the ABID reflex test and order the weak D reflex test (or other patient antigen typing test). Cancel the ABID, when it is not required for further investigation. | Occasional |
| Patient History Report65013 | The "Expired Task Processed" exception created by resulting an expired test does not display on the Patient History Report. | View the Exception Report to see all "Expired Task Processed" exceptions. | Occasional |
| Patient History Report65352 | Patient History Report does not display VistA-converted Special Instructions (SI) and Transfusion Requirements (TR). | All SIs and TRs are displayed in the Patient Search window. Note the values before leaving the Patient Search window to launch the Patient History Report.Do not create duplicate VBECS SIs and TRs for the VistA-converted ones. | Remote |
| Patient History Report65057 | An extra audit may appear on the Patient History if the patient's middle name is missing or removed and the patient's record is edited in VistA. | The patient update event is logged and displays on the Patient History Report and the Transfusion Requirements Report. | Occasional |
| Patient Testing64239 | When entering test results for a patient with a rack that has not been QC'd, the system will generate an exception "QC not performed on rack used for testing". This section will not be saved when the selected patient has a middle name or initial. | Standard practice is not to override the QC not performed warning message. Verify the QC was performed daily by reviewing the Testing Worklist Report. | Occasional |
| Patient Testing: ABO/Rh Test70151 | If a patient has no previous ABO/Rh test record (NR) and a TAS and reflex Repeat ABO/Rh are ordered, completing the Repeat ABO/Rh test first will cause the TAS to be unsavable. When attempting to save the TAS the following message will display: "Patient's ABO/Rh was modified by another user. This window will now close and any patient's test results entered WILL NOT be saved. Please refer to patient's reports for more information." | The TAS must be completed first. If a Repeat ABO/Rh order has been completed before the TAS, it must be invalidated so that the TAS may be processed. | Remote |
| Patient Testing: ABS Test63661 | VBECS will crash when multiple patient Antigen Typing tests are selected at the same time under this condition: after the testing grids appear, clicking on the Antigen tab that is not currently active, and then canceling all reagent log selection windows. | Process one Antigen Typing test at a time. | Occasional |
| Patient Testing: ABS Test65305 | A patient ABS with a Positive interpretation may be saved as a completed test with blank reaction results. The Reaction Result cells on the Testing Worklist Report will also contain blanks. | Enter all ABS reaction results before entering the ABS interpretation. | Occasional |
| Patient Testing: Antigen Typing Test64510 | VBECS will crash upon entering a result of "H" (hemolysis) in any of the Patient Antigen Typing testing grids. | Do not enter "H". Refer to the Valid Entries legend beneath the testing grids for acceptable entries. | Occasional |
| Patient Testing: Antigen Typing Test70370 | A discrepancy override is not displayed when a patient Antigen Typing test and a Repeat Antigen Typing test for the same antigen are discrepant.The discrepancy override is displayed when two patient Antigen Typing tests for the same antigen are discrepant. | Do not process a REPEAT patient antigen typing test.  | Remote |
| Patient Testing: Antigen Typing Test75709 | Some reagent suppliers reuse lot numbers allowing for two reagents with the same lot number. If a lot number is received a second time, the previously used and expired reagent is displayed instead of the most recent. | Set the number of vials for the previously used reagent to zero to make it ineligible for selection. | Remote |
| Patient Testing: DAT Test64032 | DAT grid does not properly calculate the QC status of the Polyspecific AHG reagent when multiple lot numbers are used on the same day. | View the QC data for the day for the lot number in question from the Testing Worklist Report and re-enter the results for the test grid (or repeat the testing). | Occasional |
| Patient Testing: Record a Crossmatch84332 | When a RBC and a WB unit are crossmatched for a single patient together in one transaction the order completion messages are not sent to CPRS. | Crossmatch the RBC and WB units separately. | Remote |
| Patient Testing: Record a Crossmatch70598 | A result of “R” indicating Rouleaux is erroneously allowed in the AHG and CC phases of the Antibody Screen and XM tests.  | “R” should not be used to indicate a questionable test result in the AHG and CC phases of the Antibody and XM tests. | Remote |
| Patient Testing Worklist and Testing Worklist Reports64271 | Various reagent lot numbers are not displayed including QC kit, Reverse ABO cells, PEG, LISS, or Anti-Human Globulin. | Manually record daily QC testing and reagent lot numbers to remain compliant with regulatory requirements (hard copy or spreadsheet). Retain with Testing Worklist review records. The incompletely displayed QC lot numbers and results force the blood bank to establish a policy to record results to maintain a complete record for accreditation and regulatory compliance which has been put in place by the sites. | Frequent |
| Post-Transfusion Information64698 | When transfusing multiple units, VBECS displays a tab for each unit. VBECS may crash if the tabs/units are NOT processed from left-to-right. | Process the tabs from left-to-right to avoid the crash. If VBECS crashes, no data are saved, and the information must be re-entered. | Occasional |
| Post-Transfusion Information65074 | If the "previous month" or "next month" buttons in the calendar control are clicked to set the Transfusion Start Date or the Transfusion End Date, VBECS will present an infinite loop of error messages. | Use the keyboard to set the Transfusion Start Date and Transfusion End Date fields. Once the endless errors start you must Sign Out (not disconnect) the VBECS server to make it stop. | Occasional |
| Process Incoming Shipment65000 | When a unit is received for the second time, VBECS will allow the unit to be saved with a missing ABO/Rh. When a unit is saved without a blood type, it cannot be properly confirmed nor transfused. | Verify all unit data entered into VBECS against the unit labels before saving. To correct a unit that has been entered without a blood type, inactivate the unit, and then re-enter it correctly in Incoming Shipment. | Occasional |
| Process Outgoing Shipment64558 | The outgoing shipping invoice does not display a blood unit’s CMV Negative, Sickle Cell, and Autologous unit testing status. | Handwrite pertinent information on the printed invoice, when used to ship blood products. | Occasional |
| Select Units for a Patient65035 | If ENTER is pressed when manually entering the unit ID and product code, the Select Units form acts as if the OK button has been pressed. This can cause the form to close without assigning the unit or prompting to perform XM testing. | Scan all unit fields when possible. When entering a unit manually, press the TAB key after entering the product code, or use the Find/Search feature. | Occasional |
| Select Units for a Patient65157 | A crash occurs when selecting a unit and multiple duplicate Component Requirements are set in a division. | Inactivate the duplicate Component Requirements so that only one is active in the division. | Remote |
| Select Units for a Patient70667 | VBECS allows electronic crossmatch on a patient with a history of testing for only one specimen. | FDA guidance states that it is acceptable to have two ABO/Rh typings on the same specimen, however it is highly recommended to draw two different samples and test them. | Frequent |
| Select Units for a Patient84872 | When assigning units to a batch with multiple orders, some requiring crossmatch (i.e., RBC) and some not requiring crossmatch (e.g., FFP), all orders show as completed in CPRS before the required crossmatch has been completed. | Process component orders separately. | Remote |
| Testing Worklist Report64223 | The Preliminary Rack QC section does not display the lot number of a reagent if the lot number is the same lot number for the reagent displayed before it. | Manually record missing lot numbers on the printed report. | Improbable |
| Testing Worklist Report70581 | When patient antigen typing tests are invalidated and re-entered, the Miscellaneous QC section associates the QC performed with the invalidated test with the replacement test. | Manually document any offline QC performed associated with the replacement test. | Occasional |
| Transfusion Reaction Workup (TRW)63910 | After a unit has been added to the Implicated Unit(s) list, the unit is selectable from the list to allow for data corrections. Selecting a unit from this list can cause some field data to be lost. | To correct entries for an implicated unit, click Cancel on the screen and reopen the TRW. | Occasional |
| Unit Antigen Typing70290 | Comments are not saved if the user types a custom comment instead of selecting one of the canned comments. | Click twice on the comment field until the cursor shows up in the comment drop down to enter a free text comment. Recommend entering a canned comment when possible. | Occasional |
| Unit Antigen Typing63965 | Opening a partially completed worklist causes a system crash. This can happen if database connectivity was interrupted during the previous save which causes a completed worklist to display as a partially completed worklist. This condition will also allow users to invalidate a completed worklist since it is presenting as partially completed. | Before invalidating a worklist, confirm that the worklist is incomplete by checking the testing of each unit on the worklist using the **Edit Unit Information** option. File a service desk ticket if a completed worklist is displaying as partially completed and causing a system crash. | Remote |
| Unit Antigen Typing63608 | Control cell comments are not visible after re-opening a partially completed worklist. | Testing comments are saved and can be viewed on the Testing Worklist Report.  | Occasional |
| Update Reagent Inventory70767 | When changing a lot number within update Reagent inventory, VBECS will crash if "X" is clicked to close the form with the cursor inside the Change Date control. | Type in the date rather than use the Change Date control. | Remote |
| VBECS Admin86447 | Order alerts do not work if you remove a printer and do not add it back with the same exact name, including case-sensitivity. | Navigate to Tools -> Configure Division, click on Order Alerts tab and select newly defined printer in the Printer Name dropdown. Click OK to save. | Occasional |

### Revision History

| **Date** | **Revision** | **Description** | **Author** |
| --- | --- | --- | --- |
| 9/9/21 | 1.0 | VBECS 2.3.3Modified VBECS 2.3.2 Known Defects and Anomalies, Version 4.0 to create the VBECS 2.3.3 Known Defects and Anomalies, Version 1.0. (Task 61758)Global: Replaced RTC work items with Azure DevOps work items.Changes for VBECS 2.3.3: Replaced Changes for VBECS 2.3.2 sections.Related Manuals and Materials: Updated documents to 2.3.3.**Added:**Incoming Shipment 64944Patient Testing: Antigen Typing Test 75709Select Units for a Patient 70667**Modified Likelihood of Occurrence:**ABO/Rh Confirmation 70543: Remote to OccasionalCost Accounting Report 70841: Occasional to ImprobableDiscard or Quarantine Units 70223: Improbable to OccasionalTesting Worklist Report 64223: Frequent to Improbable | BBM Team |
| 11/4/21 | 2.0 | VBECS 2.3.3Updates from IOC Test (Task 84516)Changes for VBECS 2.3.3: Added rows for CPRS 64994, Patient Testing: Record a Crossmatch 84332, and Select Units for a Patient 84872.**Added:**CPRS 64994Patient Testing: Record a Crossmatch 84332 Select Units for a Patient 84872**Modified VBECS Option:**Patient Testing: Record a Crossmatch 70598: Changed from Patient Testing: XM Test | BBM Team |
| 4/11/22 | 3.0 | VBECS 2.3.3 Rev B (Task 87272)**Added:**VBECS Admin 86447 | BBM Team |

This is the last page of *VBECS 2.3.3 Known Defects and Anomalies*.