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[bookmark: _Toc183198743]Introduction
[bookmark: OLE_LINK41][bookmark: OLE_LINK42]VistA Blood Establishment Computer Software (VBECS) 2.4.0 Rev B Release Notes contains information for changes and corrections made to VBECS in the 2.4.0 Rev B patch.  
Perform a local evaluation and risk assessment of the changes to determine the requirements for local validation of the changes, including documentation of the assessment activities. See Validation Planning.
[bookmark: _Toc183198744]Changed Features and Functions
[bookmark: _Hlk149144237]This database-only revision letter patch release is centrally focused around providing additional blood product updates released from International Council for Commonality in Blood Banking Automation, Inc. (ICCBBA) and other database-only changes to resolve known defects and requests. Blood products contained in ICCBBA releases 7.69.0 through 7.75.4 are included.
Regulatory compliance updates
· Blood product updates contained in ICCBBA releases 7.69.0 through 7.75.4.
· Minor antigen name updates to match AABB Technical Manual, 21st Edition.
Solutions to known requests and defects
· Fix known requests and defects related to the CPRS Blood Bank Report
· Fix known requests and defects related to the Patient History Report.
· Fix known requests and defects related to Process Outgoing Shipment.
· Fix known requests and defects related to system performance.
· Fix known requests and defects related to database parameter tables.
System maintenance updates	
· Clean up permissions for the VBECS SharePoint site.
· Ungroup the VBECS icons on the Windows task bar.
Table 1: Updates by Option provides a complete list of included changes. For a full list of new product codes see Table 2: New Blood Products.
[bookmark: _Toc183198745]Untestable System-Level Maintenance Changes
· Fix known requests and defects related to database functions and administrative tables.197429, 259065
[bookmark: _Toc183198746]VistA Software Dependencies
· VistA informational patch VBEC*2*15.


[bookmark: _Toc183198747]VBECS User Documents
The following is a list of all the VBECS user documents that apply to the VBECS 2.4.0 Rev B patch release. The Updated column identifies the documents that have been updated with this VBECS 2.4.0 Rev B patch release.  
These user documents are available from the VBECS SharePoint site. Refer to VBECS SharePoint Site.
	User Document
	
	
	Version
	Updated 

	VBECS 2.4.0 User Guide
	
	
	2.0
	Yes

	VBECS 2.4.0 Admin User Guide
	
	
	2.0
	Yes

	VBECS 2.4.0 Known Defects and Anomalies
	
	
	3.0
	Yes

	VBECS 2.4.0 Technical Manual-Security Guide
	
	
	2.0
	Yes

	VBECS - Echo Interface Configuration and Setup Guide
	
	
	6.0
	No

	VBECS - Erytra Interface Configuration and Setup Guide
	
	
	3.0
	No

	VBECS - ProVue Interface Configuration and Setup Guide
	
	
	6.0
	No

	VBECS - Vision Interface Configuration and Setup Guide
	
	
	5.0
	No

	VBECS Product Information
	
	
	11.0
	No





The following is a list of all the VBECS FAQ documents that apply to the VBECS 2.4.0 Rev B patch release. The Updated column identifies the documents that have been updated with this VBECS 2.4.0 Rev B patch release.  
These FAQ documents are available from the VA SharePoint site. Refer to VBECS SharePoint Site.
	FAQ Document
	Version
	Updated 

	FAQ ABO Subgroup Interpretations
	11/22/16
	No

	FAQ Accreditation Questions Regarding Laboratory Information Systems
	11/15/24
	Yes

	FAQ Ancillary VistA Validation
	11/28/17
	No

	FAQ Antibodies with No Antigen Negative Requirement
	11/25/24
	Yes

	FAQ Blood Product Table Processes
	02/21/23
	No

	FAQ CAP Comprehensive Transfusions Medicine Crossmatch Survey
	11/22/16
	No

	FAQ CPRS VBECS Order Details
	04/01/09
	No

	FAQ Database Conversion Oddballs
	09/15/09
	No

	FAQ Documenting Unhandled Exceptions
	05/21/07
	No

	FAQ Handling ABO Incompatible Transfusion Situations
	09/23/14
	No

	FAQ How to File a New Service Request for Changes to VBECS
	RETIRED
	RETIRED

	FAQ How to take faster screen shots from Remote Desktop Session
	12/17/14
	No

	FAQ Identifying a Comparable Blood Product Code
	09/03/21
	No

	FAQ Local Facilities
	06/12/20
	No

	FAQ Modification Target Not Available
	11/28/17
	No

	FAQ Multidivisional QC Rack Workaround
	11/28/17
	No

	FAQ Order Status Clarification
	11/15/24
	Yes

	FAQ Polyspecific AHG Not Used for Testing
	01/29/10
	No

	FAQ Proper Use of PIV Card with VBECS
	09/14/21
	No

	FAQ QC Setup
	01/29/10
	No

	FAQ Retesting QC
	11/15/24
	Yes

	FAQ Retrieving an Expired Order
	12/07/17
	No

	FAQ The Difference Between Transfusion Only and Full Service Facility Types
	02/08/10
	No

	FAQ Troubleshooting Remote Desktop Issues
	12/03/24
	Yes

	FAQ VBECS Blood Product Hierarchy
	03/28/19
	No

	FAQ VBECS Computer Crossmatch Decision Tree
	02/08/12
	No

	FAQ Weak D Policy
	01/11/10
	No

	FAQ What is VBECS Read-Only
	12/04/24
	Yes

	FAQ Workload Assignments Single vs Batch Processing
	07/01/20
	No




[bookmark: _Toc183198748]Customer Support
[bookmark: _Toc168887121]Contact your Local/Regional Office of Information Technology (OIT) or Laboratory Information Manager (LIM) if you encounter VistA or CPRS connection problems and for training support before contacting the Service Desk (SD).
[image: image001]Please ensure local/regional VistA Support contact information is available at all times. 
If you experience an FDA reportable adverse event (patient death or serious injury) that VBECS may have caused or contributed to, contact the Service Desk directly to enter a ticket for Blood Bank software support. 

Service Desk Contact
Call the Service Desk REDACTED
Refer to the VBECS Getting Help page on the VBECS SharePoint site for up-to-date information about contacting the Service Desk.
[bookmark: _Toc183198749]References
ISBT 128 Standard Technical Specification v 6.2.2
ISBT 128 Standard Terminology for Medical Products of Human Origin, v 7.75
United States Consensus Standard for the Uniform Labeling of Blood and Blood Components Using ISBT 128, v3.0.0
Standards for Blood Banks and Transfusion Services, 34th Edition
AABB Technical Manual, 21st Edition
 
[bookmark: _Ref183084535][bookmark: _Toc183198750]VBECS SharePoint Site
The VBECS SharePoint site provides a location for additional information related to the VBECS application such as FAQs, installation status, remote desktop and report share shortcuts, and release history.
REDACTED


[bookmark: _Toc183198751]Installation Qualification (IQ) Documentation

	                                             Test Account Upgrade to VBECS 2.4.0 Revision B

	Required Patch Installation
	VistA Patch: None
Server Patch: VBECS 2.4.0 Revision B

	Installation Process for VBECS 2.4.0 Rev B
	Installed by the BBM team, all sites will be installed in one day. The date of installation will be communicated through a ListServ announcement.

	Expected Downtime
	None

	Installation Communication for VBECS 2.4.0 Rev B
	The planned completion time for all test installs will be provided on the VBECS SharePoint site.

	Site Responsibility
	· Join VBECS-L message board on ListServ, if needed.
· Communicate any changes to your site supervisor and Point of Contact (POC) to the Pathology and Laboratory Management staff (PLM).  

	Site Record of Patch Installation
	Sites should take and save a screenshot of the VBECS Help, about window after the Test installation is performed. This displays the VBECS updated version information. 



	                                             Production Account Upgrade to VBECS 2.4.0 Revision B

	Required Patch Installation
	VistA Patch: None
Server Patch: VBECS 2.4.0 Revision B

	Installation Process for VBECS 2.4.0 Rev B
	Installed by the BBM team, all sites will be installed in one day. The date of installation will be communicated through a ListServ announcement.

	Expected Downtime
	You will be able to use VBECS during the patch installation. VBECS installation on your production account only lasts a few minutes and is invisible to users.
If you encounter any disruption or errors while viewing or printing reports during the installation, try again after a few minutes.

	Installation Communication for VBECS 2.4.0 Rev B
	The status of production installs will be provided on the VBECS SharePoint site.

	Site Responsibility
	· Perform local validation, training, and set-up requirements prior to Production installation.
· Join VBECS-L message board on ListServ, if needed.
· Communicate any changes to your site supervisor and POC to the PLM.

	Site Record of Patch Installation
	Sites should take and save a screenshot of the VBECS Help, about window after Production installation is performed. This displays the VBECS updated version information. 



[bookmark: _Ref338930387][bookmark: _Ref494275082][bookmark: _Toc183198752]Validation Planning
The following is a flowchart to help assess changes for validation planning.









[bookmark: _Ref513719669][bookmark: _Hlk521587537]
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[bookmark: _Ref100836441][bookmark: _Ref57017996][bookmark: _Toc183198753]Table 1: Updates by Option
	ID

	Option
	Problem Summary
	Change Summary
	Validation Scenario
	Change applies to my facility. (Y/N)
	Local risk Assessment 
(Low, Med, High)
	SOP revision required. If yes, identify it.
	Staff training needed. (Y/N)
	Scenarios or validation must be performed. (Y/N)

	1
	CPRS Blood Bank Report 208042
	Transfusion Reaction Workup (TRW) comments over 510 are truncated.
	TRW comments over 510 characters are no longer being truncated.
	Record a TRW and go to the Finalize TRW page.  In the Transfusion Reaction Comments section, select any Comment and enter more than 510 characters (maximum 1000 characters) in the Details section and save.  
View the TRW report to validate that the text entered in the Details section is no longer being truncated.
Go to CPRS and view the Blood Bank Report. In the Transfusion Reactions section, validate that the text entered in the Details section is no longer being truncated.
	
	
	
	
	

	2
	Patient History Report 65343
	Implicated units appear multiple times when a TRW contains multiple implicated units each associated with a comment.
	Implicated units appear once in the Transfusion Reactions section in the Patient History Report.
	Create a TRW that contains multiple implicated units each associated with a comment.  
Run the Patient History Report and view the Transfusion Reactions section to validate that the implicated units no longer display multiple times.  
	
	
	
	
	

	3
	Patient Testing Worklist
Blood Availability 258406
	Patient Testing Worklist and Blood Availability reports were taking several minutes to generate for patients with multiple automated TAS tests.
	Performance improvements to Patient Testing Worklist and Blood Availability reports for patients with multiple historical automated TAS tests.
	This scenario is applicable only to sites with performance issues with the Patient Testing Worklist and Blood Availability reports.
Prior to installation, run the Patient Testing Worklist and Blood Availability reports and record how long it takes to receive reports for a patient with a lot of historical TAS results.
After installation, run the Patient Testing Worklist and Blood Availability reports and time how long it takes to receive results.  
	
	
	
	
	

	4
	Process Outgoing Shipment / Print Invoice191657
	RBC units appear on the invoice as contaminated in certain cases when it is not contaminated
	RBC units that are not contaminated no longer appear as contaminated.
	Log in 2 PLATELET units that allow setting the =>2ml RBC Contamination flag (e.g., E3056V00)
Activate the product code if needed.
For Unit 1: check the CMV box but leave the =>2ml RBC Contamination box unchecked.
Fill in any remaining required fields and add Unit 1 to the shipment.
For Unit 2: check the CMV box AND the =>2ml RBC Contamination box.
Fill in any remaining required fields and add Unit 2 to the shipment.
Start a new Outgoing Shipment (Shipments -> Outgoing Shipment -> New Invoice).
Add Unit 1 and Unit 2 to the invoice and click Save and Yes to confirm the shipment.
Click Print Invoice to view the invoice and verify that only Unit2 indicates =>2ml RBC Contamination.
Verify both units contain Special Testing: CMV Negative, but only Unit 2 displays =>2ml RBC Contamination.
	
	
	
	
	

	5
	Throughout VBECS 170263
	Update blood product table with new ICCBBA blood product codes from 7.69.0 through 7.75.4.
	New blood products were added to VBECS.
For a full list of new product codes see Table 2.
The list of all possible non-pooled modifications for new blood products can be found on the VBECS SharePoint Site.
	Validate only blood products that you may receive from your supplier. Use Tools, Blood Products to associate the blood products with an active facility.
	
	
	
	
	

	6
	Throughout VBECS 158622
	Update antigen type names to match 21st Edition of the AABB Technical Manual.
	Minor updates to antigen type names:
· V(w) to Vw - removed brackets
· Jo(a ) to Jo(a) - removed extra spacing
· Gya to Gy(a) - added brackets
	No validation is required.
	
	
	
	
	

	7
	VBECS Database 180433
	Excess Address records exist with no corresponding Facility records.
	Address parameter table records with no corresponding Facility parameter table records have been deleted.
	No validation is required.
Note: These records are not visible to the user.
	
	
	
	
	

	8
	VBECS SharePoint Site
	SharePoint site permission changes as mandated by the OIT department
	Permissions for access to the VBECS SharePoint site have been updated as mandated by the OIT department.
	Validate that the appropriate team members have access to the VBECS SharePoint site.
	
	
	
	
	

	9
	Windows Taskbar
	The VBECS application server taskbar icons are grouped, which can hide some prompts for user input, such a PIV certificate selection.
	When applicable, dialog or message boxes will be seen as separate items on the taskbar of the VBECS server.
	No validation is required
	
	
	
	
	




[bookmark: _Ref100836461][bookmark: _Ref100836274][bookmark: _Ref100836549][bookmark: _Toc183198754][bookmark: _Ref100834973]Table 2: New Blood Products
	Component
	Product Code
	Modifiable
	Product Type
	Long Name
	Short Name
	Maximum Storage Hours

	OTHER
	EB078
	Yes
	POOLED PLATELET-RICH BUFFY COAT
	POOLED PLATELET-RICH BUFFY COAT|CPD/XX/20-24C|Irradiated|ResLeu:<1E6
	BUFFY PltRch Pool CPD IRD LUKOPR
	4

	CRYOPRECIPITATE
	EB079
	Yes
	Thawed POOLED CRYOPRECIPITATE
	Thawed POOLED CRYOPRECIPITATE|NS/XX/rt|6 units|Methylene blue-treated
	CRYO Pool Thaw NS 6U Methyl Blue
	6

	CRYOPRECIPITATE
	EB080
	Yes
	Thawed POOLED CRYOPRECIPITATE
	Thawed POOLED CRYOPRECIPITATE|NS/XX/rt|7 units|Methylene blue-treated
	CRYO Pool Thaw NS 7U Methyl Blue
	6

	CRYOPRECIPITATE
	EB081
	Yes
	Thawed POOLED CRYOPRECIPITATE
	Thawed POOLED CRYOPRECIPITATE|NS/XX/rt|8 units|Methylene blue-treated
	CRYO Pool Thaw NS 8U Methyl Blue
	6

	PLATELETS
	EB082
	No
	Frozen Apheresis PLATELETS
	Frozen Apheresis PLATELETS|ACD-A/XX/<=-18C|For mnf:noninjectable|<3E11 plts
	PLT AFR Froz ACD-A MNFNI
	8760

	PLATELETS
	EB083
	No
	Frozen Apheresis PLATELETS
	Frozen Apheresis PLATELETS|ACD-A/XX/<=-18C|For mnf:noninjectable|1st container|<3E11 plts
	PLT AFR Froz ACD-A MNFNI 1C
	8760

	PLATELETS
	EB084
	No
	Frozen Apheresis PLATELETS
	Frozen Apheresis PLATELETS|ACD-A/XX/<=-18C|For mnf:noninjectable|2nd container|<3E11 plts
	PLT AFR Froz ACD-A MNFNI 2C
	8760

	PLATELETS
	EB085
	No
	Frozen Apheresis PLATELETS
	Frozen Apheresis PLATELETS|ACD-A/XX/<=-18C|For mnf:noninjectable|3rd container|<3E11 plts
	PLT AFR Froz ACD-A MNFNI 3C
	8760

	PLATELETS
	EB086
	No
	Frozen Apheresis PLATELETS
	Frozen Apheresis PLATELETS|ACD-A/XX/<=-18C|For mnf:noninjectable|Irradiated|<3E11 plts
	PLT AFR Froz ACD-A MNFNI IRD
	8760

	PLATELETS
	EB087
	No
	Frozen Apheresis PLATELETS
	Frozen Apheresis PLATELETS|ACD-A/XX/<=-18C|For mnf:noninjectable|Irradiated|1st container|<3E11 plts
	PLT AFR Froz ACD-A MNFNI IRD 1C
	8760

	PLATELETS
	EB088
	No
	Frozen Apheresis PLATELETS
	Frozen Apheresis PLATELETS|ACD-A/XX/<=-18C|For mnf:noninjectable|Irradiated|2nd container|<3E11 plts
	PLT AFR Froz ACD-A MNFNI IRD 2C
	8760

	PLATELETS
	EB089
	No
	Frozen Apheresis PLATELETS
	Frozen Apheresis PLATELETS|ACD-A/XX/<=-18C|For mnf:noninjectable|Irradiated|3rd container|<3E11 plts
	PLT AFR Froz ACD-A MNFNI IRD 3C
	8760

	RED BLOOD CELLS
	EB090
	No
	RED BLOOD CELLS
	RED BLOOD CELLS|CP2D/XX/refg|Not for tx or mnf|LowVol:anticoag not adj
	RBC CP2D NMT LowVol
	504

	RED BLOOD CELLS
	EB091
	No
	RED BLOOD CELLS
	RED BLOOD CELLS|CP2D/XX/refg|Not for tx or mnf|ResLeu:<5E6|LowVol:anticoag not adj
	RBC CP2D NMT LUKOPR LowVol
	504

	RED BLOOD CELLS
	EB092
	No
	RED BLOOD CELLS
	RED BLOOD CELLS|CP2D>AS3/XX/refg|Not for tx or mnf|LowVol:anticoag not adj
	RBC CP2D>AS3 NMT LowVol
	1008

	RED BLOOD CELLS
	EB093
	No
	RED BLOOD CELLS
	RED BLOOD CELLS|CP2D>AS3/XX/refg|Not for tx or mnf|ResLeu:<5E6|LowVol:anticoag not adj
	RBC CP2D>AS3 NMT LUKOPR LowVol
	1008

	PLATELETS
	EB094
	No
	Frozen Apheresis PLATELETS
	Frozen Apheresis PLATELETS|ACD-A/XX/<=-18C|For mnf:noninjectable|ResLeu:<5E6|1st container|<3E11 plts
	PLT AFR Froz ACD-A MNFNI LUKOPR 1C
	8760

	PLATELETS
	EB095
	No
	Frozen Apheresis PLATELETS
	Frozen Apheresis PLATELETS|ACD-A/XX/<=-18C|For mnf:noninjectable|ResLeu:<5E6|2nd container|<3E11 plts
	PLT AFR Froz ACD-A MNFNI LUKOPR 2C
	8760

	PLATELETS
	EB096
	No
	Frozen Apheresis PLATELETS
	Frozen Apheresis PLATELETS|ACD-A/XX/<=-18C|For mnf:noninjectable|ResLeu:<5E6|3rd container|<3E11 plts
	PLT AFR Froz ACD-A MNFNI LUKOPR 3C
	8760

	PLATELETS
	EB097
	No
	Frozen Apheresis PLATELETS
	Frozen Apheresis PLATELETS|ACD-A/XX/<=-18C|For mnf:noninjectable|Irradiated|ResLeu:<5E6|1st container|<3E11 plts
	PLT AFR Froz ACD-A MNFNI IRD LUKOPR 1C
	8760

	PLATELETS
	EB098
	No
	Frozen Apheresis PLATELETS
	Frozen Apheresis PLATELETS|ACD-A/XX/<=-18C|For mnf:noninjectable|Irradiated|ResLeu:<5E6|2nd container|<3E11 plts
	PLT AFR Froz ACD-A MNFNI IRD LUKOPR 2C
	8760

	PLATELETS
	EB099
	No
	Frozen Apheresis PLATELETS
	Frozen Apheresis PLATELETS|ACD-A/XX/<=-18C|For mnf:noninjectable|Irradiated|ResLeu:<5E6|3rd container|<3E11 plts
	PLT AFR Froz ACD-A MNFNI IRD LUKOPR 3C
	8760

	RED BLOOD CELLS
	EB100
	Yes
	RED BLOOD CELLS
	RED BLOOD CELLS|CPD>AS5/500mL/refg|RBC irradiated|ResLeu:<5E6|Supernat rem/Plasma added
	RBC CPD>AS5 RBC IRD LUKOPR SR/PA
	672

	RED BLOOD CELLS
	EB101
	Yes
	Washed RED BLOOD CELLS
	Washed RED BLOOD CELLS|None/XX/refg|Open|RBC irradiated|ResLeu:<5E6|Plasma added
	RBC Wash None OPN RBC IRD LUKOPR PA
	24

	WHOLE BLOOD
	EB102
	Yes
	WHOLE BLOOD
	WHOLE BLOOD|CPD/200mL/refg
	WB CPD
	504

	WHOLE BLOOD
	EB103
	Yes
	WHOLE BLOOD
	WHOLE BLOOD|CPD/400mL/refg
	WB CPD
	504

	RED BLOOD CELLS
	EB104
	Yes
	RED BLOOD CELLS
	RED BLOOD CELLS|CPD>MAP/200mL/refg
	RBC CPD>MAP
	1008

	RED BLOOD CELLS
	EB105
	Yes
	RED BLOOD CELLS
	RED BLOOD CELLS|CPD>MAP/200mL/refg|ResLeu:<5E6
	RBC CPD>MAP LUKOPR
	1008

	RED BLOOD CELLS
	EB106
	Yes
	RED BLOOD CELLS
	RED BLOOD CELLS|CPD>MAP/200mL/refg|Irradiated|ResLeu:<5E6
	RBC CPD>MAP IRD LUKOPR
	672

	RED BLOOD CELLS
	EB107
	Yes
	RED BLOOD CELLS
	RED BLOOD CELLS|CPD>MAP/400mL/refg
	RBC CPD>MAP
	1008

	RED BLOOD CELLS
	EB108
	Yes
	RED BLOOD CELLS
	RED BLOOD CELLS|CPD>MAP/400mL/refg|ResLeu:<5E6
	RBC CPD>MAP LUKOPR
	1008

	RED BLOOD CELLS
	EB109
	Yes
	RED BLOOD CELLS
	RED BLOOD CELLS|CPD>MAP/400mL/refg|Irradiated|ResLeu:<5E6
	RBC CPD>MAP IRD LUKOPR
	672

	FRESH FROZEN PLASMA
	EB110
	Yes
	FRESH FROZEN PLASMA
	FRESH FROZEN PLASMA|CPD/XX/<=-20C|ResLeu:<5E6
	FFP CPD LUKOPR
	8760

	FRESH FROZEN PLASMA
	EB111
	Yes
	Apheresis FRESH FROZEN PLASMA
	Apheresis FRESH FROZEN PLASMA|ACD-A/XX/<=-20C|ResLeu:<5E6
	FFP AFR ACD-A LUKOPR
	8760

	PLATELETS
	EB112
	Yes
	POOLED PLATELETS
	POOLED PLATELETS|CP2D/XX/20-24C
	PLT Pool CP2D
	120

	PLATELETS
	EB113
	Yes
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A/XX/20-24C|Open|Irradiated|ResLeu:<5E6|1st container|Approx 300 E9 plts
	PLT AFR ACD-A OPN IRD LUKOPR 1C
	4

	RED BLOOD CELLS
	EB114
	No
	RED BLOOD CELLS
	RED BLOOD CELLS|CPD>SAGM/500mL/refg|ResLeu:<5E6|Supernat rem/SD-Plasma added
	RBC CPD>SAGM LUKOPR SR/SD-PA
	1008

	RED BLOOD CELLS
	EB115
	Yes
	Washed RED BLOOD CELLS
	Washed RED BLOOD CELLS|None/XX/refg|Open|Irradiated|ResLeu:<5E6|SD-Plasma added
	RBC Wash None OPN IRD LUKOPR SD-PA
	24

	RED BLOOD CELLS
	EB116
	No
	RED BLOOD CELLS
	RED BLOOD CELLS|SAGM/XX/refg|Open|ResLeu:<5E6|Supernat rem/SD-Plasma added
	RBC SAGM OPN LUKOPR SR/SD-PA
	24

	RED BLOOD CELLS
	EB117
	No
	RED BLOOD CELLS
	RED BLOOD CELLS|SAGM/XX/refg|Open|Irradiated|ResLeu:<5E6|Supernat rem/SD-Plasma added
	RBC SAGM OPN IRD LUKOPR SR/SD-PA
	24

	RED BLOOD CELLS
	EB118
	Yes
	RED BLOOD CELLS
	RED BLOOD CELLS|CPD/XX/refg|Irradiated|ResLeu:<5E6|Hct 0.75-0.9
	RBC CPD IRD LUKOPR Hct 0.75-0.9
	504

	PLATELETS
	EB119
	No
	Washed Apheresis PLATELETS
	Washed Apheresis PLATELETS|PAS-E/XX/20-24C|ResLeu:<5E6
	PLT AFR Wash PAS-E LUKOPR
	4

	PLATELETS
	EB120
	No
	Washed Apheresis PLATELETS
	Washed Apheresis PLATELETS|PAS-E/XX/20-24C|Irradiated|ResLeu:<5E6
	PLT AFR Wash PAS-E IRD LUKOPR
	4

	PLATELETS
	EB121
	Yes
	Thawed Apheresis PLATELETS
	Thawed Apheresis PLATELETS|DMSO>PAS-E/XX/20-24C|Irradiated|ResLeu:<5E6
	PLT AFR Thaw DMSO>PAS-E IRD LUKOPR
	4

	RED BLOOD CELLS
	EB122
	Yes
	RED BLOOD CELLS
	RED BLOOD CELLS|CP2D>AS3/500mL/refg|Open|RBC irradiated|ResLeu:<5E6|Supernat rem/Plasma added
	RBC CP2D>AS3 OPN RBC IRD LUKOPR SR/PA
	24

	RED BLOOD CELLS
	EB123
	Yes
	RED BLOOD CELLS
	RED BLOOD CELLS|CPD>AS5/500mL/refg|Open|RBC irradiated|ResLeu:<5E6|Supernat rem/Plasma added
	RBC CPD>AS5 OPN RBC IRD LUKOPR SR/PA
	24

	WHOLE BLOOD
	EB124
	Yes
	WHOLE BLOOD
	WHOLE BLOOD|NS/XX/rt
	WB NS
	24

	PLATELETS
	EB125
	No
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A/XX/20-24C|Not for tx or mnf|<3E11 plts
	PLT AFR ACD-A NMT
	120

	PLATELETS
	EB126
	No
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A/XX/20-24C|Not for tx or mnf|1st container|<3E11 plts
	PLT AFR ACD-A NMT 1C
	120

	PLATELETS
	EB127
	No
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A/XX/20-24C|Not for tx or mnf|2nd container|<3E11 plts
	PLT AFR ACD-A NMT 2C
	120

	PLATELETS
	EB128
	No
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A/XX/20-24C|Not for tx or mnf|3rd container|<3E11 plts
	PLT AFR ACD-A NMT 3C
	120

	PLATELETS
	EB129
	No
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A/XX/20-24C|Not for tx or mnf|ResLeu:<5E6|1st container|<3E11 plts
	PLT AFR ACD-A NMT LUKOPR 1C
	120

	PLATELETS
	EB130
	No
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A/XX/20-24C|Not for tx or mnf|ResLeu:<5E6|2nd container|<3E11 plts
	PLT AFR ACD-A NMT LUKOPR 2C
	120

	PLATELETS
	EB131
	No
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A/XX/20-24C|Not for tx or mnf|ResLeu:<5E6|3rd container|<3E11 plts
	PLT AFR ACD-A NMT LUKOPR 3C
	120

	PLATELETS
	EB132
	No
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A/XX/20-24C|Not for tx or mnf|Psoralen-treated
	PLT AFR ACD-A NMT Psoraln
	120

	PLATELETS
	EB133
	No
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A/XX/20-24C|Not for tx or mnf|1st container|Psoralen-treated
	PLT AFR ACD-A NMT 1C Psoraln
	120

	PLATELETS
	EB134
	No
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A/XX/20-24C|Not for tx or mnf|2nd container|Psoralen-treated
	PLT AFR ACD-A NMT 2C Psoraln
	120

	PLATELETS
	EB135
	No
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A/XX/20-24C|Not for tx or mnf|3rd container|Psoralen-treated
	PLT AFR ACD-A NMT 3C Psoraln
	120

	PLATELETS
	EB136
	Yes
	POOLED PLATELETS
	POOLED PLATELETS|CPD/XX/20-24C|Open|ResLeu:<5E6|4 units
	PLT Pool CPD OPN LUKOPR 4U
	4

	PLATELETS
	EB137
	Yes
	POOLED PLATELETS
	POOLED PLATELETS|CPD/XX/20-24C|ResLeu:<5E6|Plasma reduced|4 units
	PLT Pool CPD LUKOPR PLS Red 4U
	120

	PLATELETS
	EB138
	Yes
	POOLED PLATELETS
	POOLED PLATELETS|CPD/XX/20-24C|Open|ResLeu:<5E6|Plasma reduced|4 units
	PLT Pool CPD OPN LUKOPR PLS Red 4U
	4

	PLATELETS
	EB139
	Yes
	POOLED PLATELETS
	POOLED PLATELETS|CPD/XX/20-24C|Open|ResLeu:<5E6|5 units
	PLT Pool CPD OPN LUKOPR 5U
	4

	PLATELETS
	EB140
	Yes
	POOLED PLATELETS
	POOLED PLATELETS|CPD/XX/20-24C|ResLeu:<5E6|Plasma reduced|5 units
	PLT Pool CPD LUKOPR PLS Red 5U
	120

	PLATELETS
	EB141
	Yes
	POOLED PLATELETS
	POOLED PLATELETS|CPD/XX/20-24C|Open|ResLeu:<5E6|Plasma reduced|5 units
	PLT Pool CPD OPN LUKOPR PLS Red 5U
	4

	PLATELETS
	EB142
	Yes
	POOLED PLATELETS
	POOLED PLATELETS|CPD/XX/20-24C|Open|ResLeu:<5E6|6 units
	PLT Pool CPD OPN LUKOPR 6U
	4

	PLATELETS
	EB143
	Yes
	POOLED PLATELETS
	POOLED PLATELETS|CPD/XX/20-24C|ResLeu:<5E6|Plasma reduced|6 units
	PLT Pool CPD LUKOPR PLS Red 6U
	120

	PLATELETS
	EB144
	Yes
	POOLED PLATELETS
	POOLED PLATELETS|CPD/XX/20-24C|Open|ResLeu:<5E6|Plasma reduced|6 units
	PLT Pool CPD OPN LUKOPR PLS Red 6U
	4

	PLATELETS
	EB145
	Yes
	POOLED PLATELETS
	POOLED PLATELETS|CPD/XX/20-24C|Open|Irradiated|ResLeu:<5E6|4 units
	PLT Pool CPD OPN IRD LUKOPR 4U
	4

	PLATELETS
	EB146
	Yes
	POOLED PLATELETS
	POOLED PLATELETS|CPD/XX/20-24C|Irradiated|ResLeu:<5E6|Plasma reduced|4 units
	PLT Pool CPD IRD LUKOPR PLS Red 4U
	120

	PLATELETS
	EB147
	Yes
	POOLED PLATELETS
	POOLED PLATELETS|CPD/XX/20-24C|Open|Irradiated|ResLeu:<5E6|Plasma reduced|4 units
	PLT Pool CPD OPN IRD LUKOPR PLS Red 4U
	4

	PLATELETS
	EB148
	Yes
	POOLED PLATELETS
	POOLED PLATELETS|CPD/XX/20-24C|Open|Irradiated|ResLeu:<5E6|5 units
	PLT Pool CPD OPN IRD LUKOPR 5U
	4

	PLATELETS
	EB149
	Yes
	POOLED PLATELETS
	POOLED PLATELETS|CPD/XX/20-24C|Irradiated|ResLeu:<5E6|Plasma reduced|5 units
	PLT Pool CPD IRD LUKOPR PLS Red 5U
	120

	PLATELETS
	EB150
	Yes
	POOLED PLATELETS
	POOLED PLATELETS|CPD/XX/20-24C|Open|Irradiated|ResLeu:<5E6|6 units
	PLT Pool CPD OPN IRD LUKOPR 6U
	4

	PLATELETS
	EB151
	Yes
	POOLED PLATELETS
	POOLED PLATELETS|CPD/XX/20-24C|Irradiated|ResLeu:<5E6|Plasma reduced|6 units
	PLT Pool CPD IRD LUKOPR PLS Red 6U
	120

	PLATELETS
	EB152
	Yes
	POOLED PLATELETS
	POOLED PLATELETS|CPD/XX/20-24C|Open|ResLeu:<5E6|<3E11 plts
	PLT Pool CPD OPN LUKOPR
	4

	PLATELETS
	EB153
	Yes
	POOLED PLATELETS
	POOLED PLATELETS|CPD/XX/20-24C|ResLeu:<5E6|Plasma reduced|<3E11 plts
	PLT Pool CPD LUKOPR PLS Red
	120

	PLATELETS
	EB154
	Yes
	POOLED PLATELETS
	POOLED PLATELETS|CPD/XX/20-24C|Open|ResLeu:<5E6|Plasma reduced|<3E11 plts
	PLT Pool CPD OPN LUKOPR PLS Red
	4

	PLATELETS
	EB155
	Yes
	POOLED PLATELETS
	POOLED PLATELETS|CPD/XX/20-24C|Open|Irradiated|ResLeu:<5E6|<3E11 plts
	PLT Pool CPD OPN IRD LUKOPR
	4

	PLATELETS
	EB156
	Yes
	POOLED PLATELETS
	POOLED PLATELETS|CPD/XX/20-24C|Irradiated|ResLeu:<5E6|Plasma reduced|<3E11 plts
	PLT Pool CPD IRD LUKOPR PLS Red
	120

	PLATELETS
	EB157
	Yes
	POOLED PLATELETS
	POOLED PLATELETS|CPD/XX/20-24C|Open|Irradiated|ResLeu:<5E6|Plasma reduced|<3E11 plts
	PLT Pool CPD OPN IRD LUKOPR PLS Red
	4

	RED BLOOD CELLS
	EB158
	Yes
	Washed Apheresis RED BLOOD CELLS
	Washed Apheresis RED BLOOD CELLS|None/XX/refg|Open|RBC irradiated|ResLeu:<5E6|Plasma added
	RBC AFR Wash None OPN RBC IRD LUKOPR PA
	24

	RED BLOOD CELLS
	EB159
	Yes
	Washed Apheresis RED BLOOD CELLS
	Washed Apheresis RED BLOOD CELLS|None/XX/refg|Open|RBC irradiated|ResLeu:<5E6|Plasma added|1st container
	RBC AFR Wash None OPN RBC IRD LUKOPR PA 1C
	24

	RED BLOOD CELLS
	EB160
	Yes
	Washed Apheresis RED BLOOD CELLS
	Washed Apheresis RED BLOOD CELLS|None/XX/refg|Open|RBC irradiated|ResLeu:<5E6|Plasma added|2nd container
	RBC AFR Wash None OPN RBC IRD LUKOPR PA 2C
	24

	RED BLOOD CELLS
	EB161
	Yes
	RED BLOOD CELLS
	RED BLOOD CELLS|CP2D>AS3/500mL/refg|ResLeu:<5E6|O2/CO2 reduced
	RBC CP2D>AS3 LUKOPR O2/CO2 Red
	1008

	RED BLOOD CELLS
	EB162
	Yes
	RED BLOOD CELLS
	RED BLOOD CELLS|CP2D>AS3/500mL/refg|Irradiated|ResLeu:<5E6|O2/CO2 reduced
	RBC CP2D>AS3 IRD LUKOPR O2/CO2 Red
	672

	RED BLOOD CELLS
	EB163
	Yes
	Deglycerolized RED BLOOD CELLS
	Deglycerolized RED BLOOD CELLS|NS/XX/refg|Open|ResLeu:<5E6|Supernat rem|Hct 0.7-0.8
	RBC Deg NS OPN LUKOPR SR Hct 0.7-0.8
	24

	PLATELETS
	EB164
	Yes
	POOLED PLATELETS
	POOLED PLATELETS|CPD>PAS-E/XX/20-24C|ResLeu:<5E6|Supernat reduced|Buffy coat plts prep|Psoralen-treated
	PLT Pool CPD>PAS-E LUKOPR SUPRed BUFCT Psoraln
	120

	PLATELETS
	EB165
	No
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A>PAS-E/XX/20-24C|ResLeu:<5E6|Supernat reduced|Psoralen-treated
	PLT AFR ACD-A>PAS-E LUKOPR SUPRed Psoraln
	120

	PLATELETS
	EB166
	No
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A>PAS-E/XX/20-24C|ResLeu:<5E6|Supernat reduced|1st container|Psoralen-treated
	PLT AFR ACD-A>PAS-E LUKOPR SUPRed 1C Psoraln
	120

	PLATELETS
	EB167
	No
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A>PAS-E/XX/20-24C|ResLeu:<5E6|Supernat reduced|2nd container|Psoralen-treated
	PLT AFR ACD-A>PAS-E LUKOPR SUPRed 2C Psoraln
	120

	PLATELETS
	EB168
	No
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A>PAS-E/XX/20-24C|ResLeu:<5E6|Supernat reduced|3rd container|Psoralen-treated
	PLT AFR ACD-A>PAS-E LUKOPR SUPRed 3C Psoraln
	120

	PLATELETS
	EB169
	Yes
	POOLED PLATELETS
	POOLED PLATELETS|CPD>PAS-E/XX/20-24C|ResLeu:<5E6|Supernat rem/Albumin added|Buffy coat plts prep|Psoralen-treated
	PLT Pool CPD>PAS-E LUKOPR SR/ALB Add BUFCT Psoraln
	120

	PLATELETS
	EB170
	No
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A>PAS-E/XX/20-24C|ResLeu:<5E6|Supernat rem/Albumin added|Psoralen-treated
	PLT AFR ACD-A>PAS-E LUKOPR SR/ALB Add Psoraln
	120

	PLATELETS
	EB171
	No
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A>PAS-E/XX/20-24C|ResLeu:<5E6|Supernat rem/Albumin added|1st container|Psoralen-treated
	PLT AFR ACD-A>PAS-E LUKOPR SR/ALB Add 1C Psoraln
	120

	PLATELETS
	EB172
	No
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A>PAS-E/XX/20-24C|ResLeu:<5E6|Supernat rem/Albumin added|2nd container|Psoralen-treated
	PLT AFR ACD-A>PAS-E LUKOPR SR/ALB Add 2C Psoraln
	120

	PLATELETS
	EB173
	No
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A>PAS-E/XX/20-24C|ResLeu:<5E6|Supernat rem/Albumin added|3rd container|Psoralen-treated
	PLT AFR ACD-A>PAS-E LUKOPR SR/ALB Add 3C Psoraln
	120

	PLATELETS
	EB174
	Yes
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A/XX/20-24C|Open|ResLeu:<5E6|Plasma reduced/Saline added|<3E11 plts
	PLT AFR ACD-A OPN LUKOPR PLS Red/SAL Add
	4

	PLATELETS
	EB175
	Yes
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A/XX/20-24C|Open|ResLeu:<5E6|Plasma reduced/Saline added|1st container|<3E11 plts
	PLT AFR ACD-A OPN LUKOPR PLS Red/SAL Add 1C
	4

	PLATELETS
	EB176
	Yes
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A/XX/20-24C|Open|ResLeu:<5E6|Plasma reduced/Saline added|2nd container|<3E11 plts
	PLT AFR ACD-A OPN LUKOPR PLS Red/SAL Add 2C
	4

	PLATELETS
	EB177
	Yes
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A/XX/20-24C|Open|ResLeu:<5E6|Plasma reduced/Saline added|3rd container|<3E11 plts
	PLT AFR ACD-A OPN LUKOPR PLS Red/SAL Add 3C
	4

	PLATELETS
	EB178
	Yes
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A/XX/20-24C|Open|Irradiated|ResLeu:<5E6|Plasma reduced/Saline added|<3E11 plts
	PLT AFR ACD-A OPN IRD LUKOPR PLS Red/SAL Add
	4

	PLATELETS
	EB179
	Yes
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A/XX/20-24C|Open|Irradiated|ResLeu:<5E6|Plasma reduced/Saline added|1st container|<3E11 plts
	PLT AFR ACD-A OPN IRD LUKOPR PLS Red/SAL Add 1C
	4

	PLATELETS
	EB180
	Yes
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A/XX/20-24C|Open|Irradiated|ResLeu:<5E6|Plasma reduced/Saline added|3rd container|<3E11 plts
	PLT AFR ACD-A OPN IRD LUKOPR PLS Red/SAL Add 3C
	4

	PLATELETS
	EB181
	Yes
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A/XX/20-24C|Open|ResLeu:<5E6|Plasma reduced|3rd container|<3E11 plts|Psoralen-treated
	PLT AFR ACD-A OPN LUKOPR PLS Red 3C Psoraln
	4

	PLATELETS
	EB182
	Yes
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A/XX/20-24C|Open|ResLeu:<5E6|Plasma reduced/Saline added|1st container|<3E11 plts|Psoralen-treated
	PLT AFR ACD-A OPN LUKOPR PLS Red/SAL Add 1C Psoraln
	4

	PLATELETS
	EB183
	Yes
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A/XX/20-24C|Open|ResLeu:<5E6|Plasma reduced/Saline added|2nd container|<3E11 plts|Psoralen-treated
	PLT AFR ACD-A OPN LUKOPR PLS Red/SAL Add 2C Psoraln
	4

	PLATELETS
	EB184
	Yes
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A/XX/20-24C|Open|ResLeu:<5E6|Plasma reduced/Saline added|3rd container|<3E11 plts|Psoralen-treated
	PLT AFR ACD-A OPN LUKOPR PLS Red/SAL Add 3C Psoraln
	4

	PLATELETS
	EB185
	Yes
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A/XX/20-24C|Irradiated|ResLeu:<5E6|Plasma reduced|4th container|<3E11 plts
	PLT AFR ACD-A IRD LUKOPR PLS Red 4C
	120

	PLATELETS
	EB186
	Yes
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A/XX/20-24C|Open|Irradiated|ResLeu:<5E6|Plasma reduced/Saline added|4th container|<3E11 plts
	PLT AFR ACD-A OPN IRD LUKOPR PLS Red/SAL Add 4C
	4

	WHOLE BLOOD
	EB187
	Yes
	WHOLE BLOOD
	WHOLE BLOOD|CPD/450mL/rt|ResLeu:<1E6
	WB CPD LUKOPR
	504

	WHOLE BLOOD
	EB188
	Yes
	WHOLE BLOOD
	WHOLE BLOOD|CPD/450mL/rt|Irradiated|ResLeu:<1E6
	WB CPD IRD LUKOPR
	504

	PLATELETS
	EB189
	No
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A>PAS-E/XX/20-24C|ResLeu:<1E6|1st container|Psoralen-treated
	PLT AFR ACD-A>PAS-E LUKOPR 1C Psoraln
	120

	PLATELETS
	EB190
	No
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A>PAS-E/XX/20-24C|ResLeu:<1E6|2nd container|Psoralen-treated
	PLT AFR ACD-A>PAS-E LUKOPR 2C Psoraln
	120

	PLATELETS
	EB191
	No
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A>PAS-E/XX/20-24C|ResLeu:<1E6|Supernat reduced|1st container|Psoralen-treated
	PLT AFR ACD-A>PAS-E LUKOPR SUPRed 1C Psoraln
	120

	PLATELETS
	EB192
	No
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A>PAS-E/XX/20-24C|ResLeu:<1E6|Supernat reduced|2nd container|Psoralen-treated
	PLT AFR ACD-A>PAS-E LUKOPR SUPRed 2C Psoraln
	120

	PLATELETS
	EB193
	No
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A>PAS-E/XX/20-24C|ResLeu:<1E6|1st container|Approx 120 E9 plts|Psoralen-treated
	PLT AFR ACD-A>PAS-E LUKOPR 1C Psoraln
	120

	PLATELETS
	EB194
	No
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A>PAS-E/XX/20-24C|ResLeu:<1E6|2nd container|Approx 120 E9 plts|Psoralen-treated
	PLT AFR ACD-A>PAS-E LUKOPR 2C Psoraln
	120

	FRESH FROZEN PLASMA
	EB195
	Yes
	Apheresis FRESH FROZEN PLASMA
	Apheresis FRESH FROZEN PLASMA|ACD-A/XX/<=-25C|ResLeu:<1E6|>=200mL<400mL|Frozen <=6h|1st container
	FFP AFR ACD-A LUKOPR 1C
	8760

	FRESH FROZEN PLASMA
	EB196
	Yes
	Apheresis FRESH FROZEN PLASMA
	Apheresis FRESH FROZEN PLASMA|ACD-A/XX/<=-25C|ResLeu:<1E6|>=200mL<400mL|Frozen <=6h|2nd container
	FFP AFR ACD-A LUKOPR 2C
	8760

	FRESH FROZEN PLASMA
	EB197
	Yes
	Apheresis FRESH FROZEN PLASMA
	Apheresis FRESH FROZEN PLASMA|ACD-A/XX/<=-25C|ResLeu:<1E6|>=200mL<400mL|Frozen <=6h|3rd container
	FFP AFR ACD-A LUKOPR 3C
	8760

	FRESH FROZEN PLASMA
	EB198
	Yes
	Apheresis FRESH FROZEN PLASMA
	Apheresis FRESH FROZEN PLASMA|ACD-A/XX/<=-25C|ResLeu:<1E6|<125 mL|Frozen <=6h
	FFP AFR ACD-A LUKOPR
	8760

	FRESH FROZEN PLASMA
	EB199
	No
	Apheresis FRESH FROZEN PLASMA
	Apheresis FRESH FROZEN PLASMA|ACD-A/XX/<=-25C|ResLeu:<1E6|<125 mL|Frozen <=6h|Quar:>=4m/retested
	FFP AFR ACD-A LUKOPR QUAR
	8760

	FRESH FROZEN PLASMA
	EB200
	No
	Thawed Apheresis FRESH FROZEN PLASMA
	Thawed Apheresis FRESH FROZEN PLASMA|ACD-A/XX/rt|ResLeu:<1E6|<125 mL|Frozen <=6h|Quar:>=4m/retested
	FFP AFR Thaw ACD-A LUKOPR QUAR
	24

	FRESH FROZEN PLASMA
	EB201
	No
	Apheresis FRESH FROZEN PLASMA
	Apheresis FRESH FROZEN PLASMA|ACD-A/XX/<=-25C|ResLeu:<1E6|>=200mL<400mL|Frozen <=6h|1st container|Quar:>=4m/retested
	FFP AFR ACD-A LUKOPR 1C QUAR
	8760

	FRESH FROZEN PLASMA
	EB202
	No
	Apheresis FRESH FROZEN PLASMA
	Apheresis FRESH FROZEN PLASMA|ACD-A/XX/<=-25C|ResLeu:<1E6|>=200mL<400mL|Frozen <=6h|2nd container|Quar:>=4m/retested
	FFP AFR ACD-A LUKOPR 2C QUAR
	8760

	FRESH FROZEN PLASMA
	EB203
	No
	Apheresis FRESH FROZEN PLASMA
	Apheresis FRESH FROZEN PLASMA|ACD-A/XX/<=-25C|ResLeu:<1E6|>=200mL<400mL|Frozen <=6h|3rd container|Quar:>=4m/retested
	FFP AFR ACD-A LUKOPR 3C QUAR
	8760

	FRESH FROZEN PLASMA
	EB204
	No
	Thawed Apheresis FRESH FROZEN PLASMA
	Thawed Apheresis FRESH FROZEN PLASMA|ACD-A/XX/rt|ResLeu:<1E6|>=200mL<400mL|Frozen <=6h|1st container|Quar:>=4m/retested
	FFP AFR Thaw ACD-A LUKOPR 1C QUAR
	24

	FRESH FROZEN PLASMA
	EB205
	No
	Thawed Apheresis FRESH FROZEN PLASMA
	Thawed Apheresis FRESH FROZEN PLASMA|ACD-A/XX/rt|ResLeu:<1E6|>=200mL<400mL|Frozen <=6h|2nd container|Quar:>=4m/retested
	FFP AFR Thaw ACD-A LUKOPR 2C QUAR
	24

	FRESH FROZEN PLASMA
	EB206
	No
	Thawed Apheresis FRESH FROZEN PLASMA
	Thawed Apheresis FRESH FROZEN PLASMA|ACD-A/XX/rt|ResLeu:<1E6|>=200mL<400mL|Frozen <=6h|3rd container|Quar:>=4m/retested
	FFP AFR Thaw ACD-A LUKOPR 3C QUAR
	24

	PLATELETS
	EB207
	Yes
	POOLED PLATELETS
	POOLED PLATELETS|CPD>PAS-E/XX/20-24C|ResLeu:<1E6|Buffy coat plts prep|Psoralen-treated
	PLT Pool CPD>PAS-E LUKOPR BUFCT Psoraln
	120

	PLATELETS
	EB208
	Yes
	POOLED PLATELETS
	POOLED PLATELETS|CPD>PAS-E/XX/20-24C|ResLeu:<1E6|Supernat reduced|Buffy coat plts prep|Psoralen-treated
	PLT Pool CPD>PAS-E LUKOPR SUPRed BUFCT Psoraln
	120

	PLATELETS
	EB209
	No
	Frozen POOLED PLATELETS
	Frozen POOLED PLATELETS|CPD/XX/<=-18C|For mnf:noninjectable|ResLeu:<5E6|4 units
	PLT Pool Froz CPD MNFNI LUKOPR 4U
	8760

	PLATELETS
	EB210
	No
	Frozen POOLED PLATELETS
	Frozen POOLED PLATELETS|CPD/XX/<=-18C|For mnf:noninjectable|ResLeu:<5E6|5 units
	PLT Pool Froz CPD MNFNI LUKOPR 5U
	8760

	PLATELETS
	EB211
	No
	Frozen POOLED PLATELETS
	Frozen POOLED PLATELETS|CPD/XX/<=-18C|For mnf:noninjectable|ResLeu:<5E6|6 units
	PLT Pool Froz CPD MNFNI LUKOPR 6U
	8760

	PLATELETS
	EB212
	No
	Frozen POOLED PLATELETS
	Frozen POOLED PLATELETS|CPD/XX/<=-18C|For mnf:noninjectable|Irradiated|ResLeu:<5E6|4 units
	PLT Pool Froz CPD MNFNI IRD LUKOPR 4U
	8760

	PLATELETS
	EB213
	No
	Frozen POOLED PLATELETS
	Frozen POOLED PLATELETS|CPD/XX/<=-18C|For mnf:noninjectable|Irradiated|ResLeu:<5E6|5 units
	PLT Pool Froz CPD MNFNI IRD LUKOPR 5U
	8760

	PLATELETS
	EB214
	No
	Frozen POOLED PLATELETS
	Frozen POOLED PLATELETS|CPD/XX/<=-18C|For mnf:noninjectable|Irradiated|ResLeu:<5E6|6 units
	PLT Pool Froz CPD MNFNI IRD LUKOPR 6U
	8760

	RED BLOOD CELLS
	EB215
	Yes
	RED BLOOD CELLS
	RED BLOOD CELLS|CP2D>AS3/500mL/refg|RBC irradiated|ResLeu:NS|Supernat rem/Plasma added
	RBC CP2D>AS3 RBC IRD LUKOPR SR/PA
	672

	RED BLOOD CELLS
	EB216
	Yes
	Apheresis RED BLOOD CELLS
	Apheresis RED BLOOD CELLS|CP2D>AS3/XX/refg|RBC irradiated|ResLeu:NS|Supernat rem/Plasma added|1st container
	RBC AFR CP2D>AS3 RBC IRD LUKOPR SR/PA 1C
	672

	RED BLOOD CELLS
	EB217
	Yes
	Apheresis RED BLOOD CELLS
	Apheresis RED BLOOD CELLS|CP2D>AS3/XX/refg|RBC irradiated|ResLeu:NS|Supernat rem/Plasma added|2nd container
	RBC AFR CP2D>AS3 RBC IRD LUKOPR SR/PA 2C
	672

	RED BLOOD CELLS
	EB218
	Yes
	Apheresis RED BLOOD CELLS
	Apheresis RED BLOOD CELLS|ACD-A>AS1/XX/refg|RBC irradiated|ResLeu:NS|Supernat rem/Plasma added|1st container
	RBC AFR ACD-A>AS1 RBC IRD LUKOPR SR/PA 1C
	672

	RED BLOOD CELLS
	EB219
	Yes
	Apheresis RED BLOOD CELLS
	Apheresis RED BLOOD CELLS|ACD-A>AS1/XX/refg|RBC irradiated|ResLeu:NS|Supernat rem/Plasma added|2nd container
	RBC AFR ACD-A>AS1 RBC IRD LUKOPR SR/PA 2C
	672

	PLATELETS
	EB220
	No
	POOLED PLATELETS
	POOLED PLATELETS|PAS-E/XX/20-24C|ResLeu:<1E6|Supernat rem/Plasma added|Buffy coat plts prep|4 units
	PLT Pool PAS-E LUKOPR SR/PA BUFCT 4U
	120

	PLATELETS
	EB221
	No
	POOLED PLATELETS
	POOLED PLATELETS|PAS-E/XX/20-24C|ResLeu:<1E6|Supernat rem/Plasma added|Buffy coat plts prep|5 units
	PLT Pool PAS-E LUKOPR SR/PA BUFCT 5U
	120

	PLATELETS
	EB222
	No
	POOLED PLATELETS
	POOLED PLATELETS|PAS-E/XX/20-24C|ResLeu:<1E6|Supernat rem/Plasma added|Buffy coat plts prep|6 units
	PLT Pool PAS-E LUKOPR SR/PA BUFCT 6U
	120

	PLATELETS
	EB223
	Yes
	POOLED PLATELETS
	POOLED PLATELETS|CPD/XX/20-24C|ResLeu:<5E6|<3E11 plts|Bacterial monitoring >=36h
	PLT Pool CPD LUKOPR BacMtr >=36h
	120

	PLATELETS
	EB224
	Yes
	POOLED PLATELETS
	POOLED PLATELETS|CPD/XX/20-24C|Irradiated|ResLeu:<5E6|<3E11 plts|Bacterial monitoring >=36h
	PLT Pool CPD IRD LUKOPR BacMtr >=36h
	120

	WHOLE BLOOD
	EB225
	Yes
	WHOLE BLOOD
	WHOLE BLOOD|CPD/XX/refg|ResLeu:<5E6|Platelets spared
	WB CPD LUKOPR PLT S
	504

	FRESH FROZEN PLASMA
	EB226
	Yes
	Apheresis FRESH FROZEN PLASMA
	Apheresis FRESH FROZEN PLASMA|NaCitrate/XX/<=-25C|ResLeu:<1E6|3rd container
	FFP AFR NaCit LUKOPR 3C
	8760

	FRESH FROZEN PLASMA
	EB227
	Yes
	Apheresis FRESH FROZEN PLASMA
	Apheresis FRESH FROZEN PLASMA|NaCitrate/XX/<=-25C|Irradiated|ResLeu:<1E6|3rd container
	FFP AFR NaCit IRD LUKOPR 3C
	8760

	FRESH FROZEN PLASMA
	EB228
	Yes
	Thawed Apheresis FRESH FROZEN PLASMA
	Thawed Apheresis FRESH FROZEN PLASMA|NaCitrate/XX/refg|Irradiated|ResLeu:<1E6|3rd container
	FFP AFR Thaw NaCit IRD LUKOPR 3C
	24

	FRESH FROZEN PLASMA
	EB229
	Yes
	Apheresis PLASMA
	Apheresis PLASMA|ACD-A/XX/<=-18C|1st container
	PLASMA AFR ACD-A 1C
	8760

	FRESH FROZEN PLASMA
	EB230
	Yes
	Apheresis PLASMA
	Apheresis PLASMA|ACD-A/XX/<=-18C|2nd container
	PLASMA AFR ACD-A 2C
	8760

	FRESH FROZEN PLASMA
	EB231
	Yes
	Apheresis PLASMA
	Apheresis PLASMA|ACD-A/XX/<=-18C|3rd container
	PLASMA AFR ACD-A 3C
	8760

	FRESH FROZEN PLASMA
	EB232
	Yes
	Apheresis PLASMA
	Apheresis PLASMA|ACD-A/XX/<=-18C|4th container
	PLASMA AFR ACD-A 4C
	8760

	FRESH FROZEN PLASMA
	EB233
	Yes
	Liquid PLASMA
	Liquid PLASMA|CP2D/XX/refg|ResLeu:<5E6
	PLASMA Liq CP2D LUKOPR
	624

	RED BLOOD CELLS
	EB234
	Yes
	Washed RED BLOOD CELLS
	Washed RED BLOOD CELLS|None/XX/refg|Open|ResLeu:<5E6|Supernat reduced
	RBC Wash None OPN LUKOPR SUPRed
	24

	RED BLOOD CELLS
	EB235
	Yes
	Washed RED BLOOD CELLS
	Washed RED BLOOD CELLS|None/XX/refg|Open|Irradiated|ResLeu:<5E6|Supernat reduced
	RBC Wash None OPN IRD LUKOPR SUPRed
	24

	RED BLOOD CELLS
	EB236
	Yes
	Washed Apheresis RED BLOOD CELLS
	Washed Apheresis RED BLOOD CELLS|None/XX/refg|Open|ResLeu:<5E6|Supernat reduced
	RBC AFR Wash None OPN LUKOPR SUPRed
	24

	RED BLOOD CELLS
	EB237
	Yes
	Washed Apheresis RED BLOOD CELLS
	Washed Apheresis RED BLOOD CELLS|None/XX/refg|Open|ResLeu:<5E6|Supernat reduced|1st container
	RBC AFR Wash None OPN LUKOPR SUPRed 1C
	24

	RED BLOOD CELLS
	EB238
	Yes
	Washed Apheresis RED BLOOD CELLS
	Washed Apheresis RED BLOOD CELLS|None/XX/refg|Open|ResLeu:<5E6|Supernat reduced|2nd container
	RBC AFR Wash None OPN LUKOPR SUPRed 2C
	24

	RED BLOOD CELLS
	EB239
	Yes
	Washed Apheresis RED BLOOD CELLS
	Washed Apheresis RED BLOOD CELLS|None/XX/refg|Open|Irradiated|ResLeu:<5E6|Supernat reduced
	RBC AFR Wash None OPN IRD LUKOPR SUPRed
	24

	RED BLOOD CELLS
	EB240
	Yes
	Washed Apheresis RED BLOOD CELLS
	Washed Apheresis RED BLOOD CELLS|None/XX/refg|Open|Irradiated|ResLeu:<5E6|Supernat reduced|1st container
	RBC AFR Wash None OPN IRD LUKOPR SUPRed 1C
	24

	RED BLOOD CELLS
	EB241
	Yes
	Washed Apheresis RED BLOOD CELLS
	Washed Apheresis RED BLOOD CELLS|None/XX/refg|Open|Irradiated|ResLeu:<5E6|Supernat reduced|2nd container
	RBC AFR Wash None OPN IRD LUKOPR SUPRed 2C
	24

	PLATELETS
	EB242
	Yes
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A>PAS-C/XX/refg|Open|ResLeu:<5E6|Psoralen-treated
	PLT AFR ACD-A>PAS-C OPN LUKOPR Psoraln
	4

	PLATELETS
	EB243
	Yes
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A>PAS-C/XX/refg|Open|ResLeu:<5E6|1st container|Psoralen-treated
	PLT AFR ACD-A>PAS-C OPN LUKOPR 1C Psoraln
	4

	PLATELETS
	EB244
	Yes
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A>PAS-C/XX/refg|Open|ResLeu:<5E6|2nd container|Psoralen-treated
	PLT AFR ACD-A>PAS-C OPN LUKOPR 2C Psoraln
	4

	PLATELETS
	EB245
	Yes
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A>PAS-C/XX/refg|Open|ResLeu:<5E6|3rd container|Psoralen-treated
	PLT AFR ACD-A>PAS-C OPN LUKOPR 3C Psoraln
	4

	OTHER
	EB246
	Yes
	PLATELET-RICH BUFFY COAT
	PLATELET-RICH BUFFY COAT|CPDA-1/500mL/20-24C
	BUFFY PltRch CPDA-1
	24

	OTHER
	EB247
	Yes
	PLATELET-RICH BUFFY COAT
	PLATELET-RICH BUFFY COAT|CPDA-1/500mL/20-24C|Irradiated
	BUFFY PltRch CPDA-1 IRD
	24

	PLATELETS
	EB248
	Yes
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A>PAS-B/XX/refg|ResLeu:<5E6|Psoralen-treated
	PLT AFR ACD-A>PAS-B LUKOPR Psoraln
	120

	PLATELETS
	EB249
	Yes
	POOLED PLATELETS
	POOLED PLATELETS|CPD>PAS-E/XX/refg|ResLeu:<5E6|Buffy coat plts prep|Psoralen-treated
	PLT Pool CPD>PAS-E LUKOPR BUFCT Psoraln
	120

	PLATELETS
	EB250
	Yes
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A>PAS-E/XX/refg|ResLeu:<5E6|1st container
	PLT AFR ACD-A>PAS-E LUKOPR 1C
	120

	PLATELETS
	EB251
	Yes
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A>PAS-E/XX/refg|ResLeu:<5E6|2nd container
	PLT AFR ACD-A>PAS-E LUKOPR 2C
	120

	PLATELETS
	EB252
	Yes
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A>PAS-E/XX/refg|Irradiated|ResLeu:<5E6|1st container
	PLT AFR ACD-A>PAS-E IRD LUKOPR 1C
	120

	PLATELETS
	EB253
	Yes
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A>PAS-E/XX/refg|Irradiated|ResLeu:<5E6|2nd container
	PLT AFR ACD-A>PAS-E IRD LUKOPR 2C
	120

	PLATELETS
	EB254
	Yes
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A>PAS-E/XX/refg|ResLeu:<5E6|1st container|Psoralen-treated
	PLT AFR ACD-A>PAS-E LUKOPR 1C Psoraln
	120

	PLATELETS
	EB255
	Yes
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A>PAS-E/XX/refg|ResLeu:<5E6|2nd container|Psoralen-treated
	PLT AFR ACD-A>PAS-E LUKOPR 2C Psoraln
	120

	FRESH FROZEN PLASMA
	EB256
	Yes
	Apheresis PLASMA
	Apheresis PLASMA|ACD-A/XX/<=-18C|Frozen <=24h|1st container|Psoralen-treated
	PLASMA AFR ACD-A 1C Psoraln
	8760

	FRESH FROZEN PLASMA
	EB257
	Yes
	Apheresis PLASMA
	Apheresis PLASMA|ACD-A/XX/<=-18C|Frozen <=24h|2nd container|Psoralen-treated
	PLASMA AFR ACD-A 2C Psoraln
	8760

	FRESH FROZEN PLASMA
	EB258
	Yes
	Apheresis PLASMA
	Apheresis PLASMA|ACD-A/XX/<=-18C|Frozen <=24h|3rd container|Psoralen-treated
	PLASMA AFR ACD-A 3C Psoraln
	8760

	PLATELETS
	EB259
	Yes
	POOLED PLATELETS
	POOLED PLATELETS|CPD>PAS-E/XX/20-24C|ResLeu:<1E6|Supernat reduced
	PLT Pool CPD>PAS-E LUKOPR SUPRed
	120

	PLATELETS
	EB260
	Yes
	POOLED PLATELETS
	POOLED PLATELETS|CPD>PAS-E/XX/20-24C|Irradiated|ResLeu:<1E6|Supernat reduced
	PLT Pool CPD>PAS-E IRD LUKOPR SUPRed
	120

	PLATELETS
	EB261
	No
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A>PAS-E/XX/20-24C|Open|ResLeu:<5E6|Plasma reduced/Saline added|1st container|Psoralen-treated
	PLT AFR ACD-A>PAS-E OPN LUKOPR PLS Red/SAL Add 1C Psoraln
	4

	PLATELETS
	EB262
	No
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A>PAS-E/XX/20-24C|Open|ResLeu:<5E6|Plasma reduced/Saline added|2nd container|Psoralen-treated
	PLT AFR ACD-A>PAS-E OPN LUKOPR PLS Red/SAL Add 2C Psoraln
	4

	OTHER
	EB263
	Yes
	PLATELET-RICH BUFFY COAT
	PLATELET-RICH BUFFY COAT|CPD/350mL/20-24C
	BUFFY PltRch CPD
	24

	OTHER
	EB264
	Yes
	PLATELET-RICH BUFFY COAT
	PLATELET-RICH BUFFY COAT|CPD/350mL/20-24C|Irradiated
	BUFFY PltRch CPD IRD
	24

	PLATELETS
	EB265
	No
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A>PAS-E/XX/20-24C|Open|ResLeu:<5E6|Plasma reduced|1st container
	PLT AFR ACD-A>PAS-E OPN LUKOPR PLS Red 1C
	4

	PLATELETS
	EB266
	No
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A>PAS-E/XX/20-24C|Open|ResLeu:<5E6|Plasma reduced|2nd container
	PLT AFR ACD-A>PAS-E OPN LUKOPR PLS Red 2C
	4

	PLATELETS
	EB267
	No
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A>PAS-E/XX/20-24C|Open|ResLeu:<5E6|Plasma reduced|1st container|Psoralen-treated
	PLT AFR ACD-A>PAS-E OPN LUKOPR PLS Red 1C Psoraln
	4

	PLATELETS
	EB268
	No
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A>PAS-E/XX/20-24C|Open|ResLeu:<5E6|Plasma reduced|2nd container|Psoralen-treated
	PLT AFR ACD-A>PAS-E OPN LUKOPR PLS Red 2C Psoraln
	4

	PLATELETS
	EB269
	No
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A>PAS-E/XX/20-24C|Open|Irradiated|ResLeu:<5E6|Plasma reduced|1st container
	PLT AFR ACD-A>PAS-E OPN IRD LUKOPR PLS Red 1C
	4

	PLATELETS
	EB270
	No
	Apheresis PLATELETS
	Apheresis PLATELETS|ACD-A>PAS-E/XX/20-24C|Open|Irradiated|ResLeu:<5E6|Plasma reduced|2nd container
	PLT AFR ACD-A>PAS-E OPN IRD LUKOPR PLS Red 2C
	4

	FRESH FROZEN PLASMA
	EB271
	Yes
	Thawed PLASMA
	Thawed PLASMA|CPD/XX/refg|1st container
	PLASMA Thaw CPD 1C
	120

	FRESH FROZEN PLASMA
	EB272
	Yes
	Thawed PLASMA
	Thawed PLASMA|CPD/XX/refg|2nd container
	PLASMA Thaw CPD 2C
	120

	FRESH FROZEN PLASMA
	EB273
	Yes
	Thawed PLASMA
	Thawed PLASMA|CPD/XX/refg|3rd container
	PLASMA Thaw CPD 3C
	120

	PLATELETS
	EB274
	Yes
	POOLED PLATELETS
	POOLED PLATELETS|CPD>PAS-E/XX/20-24C|Open|ResLeu:<5E6|Plasma reduced/Saline added|Buffy coat plts prep|Psoralen-treated
	PLT Pool CPD>PAS-E OPN LUKOPR PLS Red/SAL Add BUFCT Psoraln
	4

	RED BLOOD CELLS
	EB275
	Yes
	RED BLOOD CELLS
	RED BLOOD CELLS|CPD>AS5/450mL/refg|ResLeu:NS
	RBC CPD>AS5 LUKOPR
	1008

	RED BLOOD CELLS
	EB276
	Yes
	RED BLOOD CELLS
	RED BLOOD CELLS|CPD>AS5/XX/refg|ResLeu:NS|LowVol:anticoag not adj
	RBC CPD>AS5 LUKOPR LowVol
	1008

	RED BLOOD CELLS
	EB277
	Yes
	RED BLOOD CELLS
	RED BLOOD CELLS|CPD>AS5/XX/refg|ResLeu:<1E6|LowVol:anticoag not adj
	RBC CPD>AS5 LUKOPR LowVol
	1008

	FRESH FROZEN PLASMA
	EB278
	Yes
	FRESH FROZEN PLASMA
	FRESH FROZEN PLASMA|CPD/450mL/<=-18C|Frozen <=8h
	FFP CPD
	8760

	FRESH FROZEN PLASMA
	EB279
	Yes
	Apheresis FRESH FROZEN PLASMA
	Apheresis FRESH FROZEN PLASMA|NaCitrate/XX/<=-25C|Frozen <=18h|Psoralen-treated
	FFP AFR NaCit Psoraln
	8760

	FRESH FROZEN PLASMA
	EB280
	Yes
	POOLED FRESH FROZEN PLASMA
	POOLED FRESH FROZEN PLASMA|CPD/XX/<=-18C|From 5 donors
	FFP Pool CPD 5DNR
	8760

	RED BLOOD CELLS
	EB281
	Yes
	RED BLOOD CELLS
	RED BLOOD CELLS|CPD>AS1/500mL/refg|Open|RBC irradiated|ResLeu:NS|Plasma added
	RBC CPD>AS1 OPN RBC IRD LUKOPR PA
	24

	RED BLOOD CELLS
	EB282
	No
	RED BLOOD CELLS
	RED BLOOD CELLS|SAGM/XX/refg|Irradiated|ResLeu:<1E6|LowVol:anticoag adj
	RBC SAGM IRD LUKOPR LowVol
	672

	FRESH FROZEN PLASMA
	X0014
	Yes
	Thawed SOLVENT DETERGENT POOLED PLASMA
	Thawed SOLVENT DETERGENT POOLED PLASMA|NS/NS/refg|ABO independent
	PLASMA Pool Thaw Det Sol NS ABO ind
	24
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