- VistA Blood Establishment Computer Software (VBECS) 2.4.0 Rev B

- Release Notes Version 2.0

- Department of Veterans Affairs

- Enterprise Project Management Office

# Revision History

| Date | Revision | Description | Author |
| --- | --- | --- | --- |
| 11/22/24 | 1.0 | VBECS 2.4.0 Rev B Initial version (Task 266824) | BBM Team |
| 12/19/24 | 2.0 | VBECS 2.4.0 Rev B Updates from IOC testing. (Task 285837) VBECS User Documents table: KDA row changed to Version 3.0. FAQ documents table: 1 row added and 6 rows updated for modified or retired FAQs. | BBM Team |

Table of Contents

Revision History	i

Introduction	1

Changed Features and Functions	1

Untestable System-Level Maintenance Changes	1

VistA Software Dependencies	1

VBECS User Documents	2

Customer Support	4

References	4

VBECS SharePoint Site	4

Installation Qualification (IQ) Documentation	5

Validation Planning	6

Table 1: Updates by Option	7

Table 2: New Blood Products	11

# Introduction

VistA Blood Establishment Computer Software (VBECS) 2.4.0 Rev B Release Notes contains information for changes and corrections made to VBECS in the 2.4.0 Rev B patch.

- Perform a local evaluation and risk assessment of the changes to determine the requirements for local validation of the changes, including documentation of the assessment activities. See Validation Planning.

# Changed Features and Functions

This database-only revision letter patch release is centrally focused around providing additional blood product updates released from International Council for Commonality in Blood Banking Automation, Inc. (ICCBBA) and other database-only changes to resolve known defects and requests. Blood products contained in ICCBBA releases 7.69.0 through 7.75.4 are included.

Regulatory compliance updates

- Blood product updates contained in ICCBBA releases 7.69.0 through 7.75.4.

- Minor antigen name updates to match AABB Technical Manual, 21st Edition.

Solutions to known requests and defects

- Fix known requests and defects related to the CPRS Blood Bank Report

- Fix known requests and defects related to the Patient History Report.

- Fix known requests and defects related to Process Outgoing Shipment.

- Fix known requests and defects related to system performance.

- Fix known requests and defects related to database parameter tables.

System maintenance updates

- Clean up permissions for the VBECS SharePoint site.

- Ungroup the VBECS icons on the Windows task bar.

- Table 1: Updates by Option provides a complete list of included changes. For a full list of new product codes see Table 2: New Blood Products.

## Untestable System-Level Maintenance Changes

- Fix known requests and defects related to database functions and administrative tables.197429, 259065

## VistA Software Dependencies

- VistA informational patch VBEC*2*15.

## VBECS User Documents

The following is a list of all the VBECS user documents that apply to the VBECS 2.4.0 Rev B patch release. The Updated column identifies the documents that have been updated with this VBECS 2.4.0 Rev B patch release.

These user documents are available from the VBECS SharePoint site. Refer to VBECS SharePoint Site.

| User Document |  |  | Version | Updated |
| --- | --- | --- | --- | --- |
| VBECS 2.4.0 User Guide |  |  | 2.0 | Yes |
| VBECS 2.4.0 Admin User Guide |  |  | 2.0 | Yes |
| VBECS 2.4.0 Known Defects and Anomalies |  |  | 3.0 | Yes |
| VBECS 2.4.0 Technical Manual-Security Guide |  |  | 2.0 | Yes |
| VBECS - Echo Interface Configuration and Setup Guide |  |  | 6.0 | No |
| VBECS - Erytra Interface Configuration and Setup Guide |  |  | 3.0 | No |
| VBECS - ProVue Interface Configuration and Setup Guide |  |  | 6.0 | No |
| VBECS - Vision Interface Configuration and Setup Guide |  |  | 5.0 | No |
| VBECS Product Information |  |  | 11.0 | No |

The following is a list of all the VBECS FAQ documents that apply to the VBECS 2.4.0 Rev B patch release. The Updated column identifies the documents that have been updated with this VBECS 2.4.0 Rev B patch release.

These FAQ documents are available from the VA SharePoint site. Refer to VBECS SharePoint Site.

| FAQ Document | Version | Updated |
| --- | --- | --- |
| FAQ ABO Subgroup Interpretations | 11/22/16 | No |
| FAQ Accreditation Questions Regarding Laboratory Information Systems | 11/15/24 | Yes |
| FAQ Ancillary VistA Validation | 11/28/17 | No |
| FAQ Antibodies with No Antigen Negative Requirement | 11/25/24 | Yes |
| FAQ Blood Product Table Processes | 02/21/23 | No |
| FAQ CAP Comprehensive Transfusions Medicine Crossmatch Survey | 11/22/16 | No |
| FAQ CPRS VBECS Order Details | 04/01/09 | No |
| FAQ Database Conversion Oddballs | 09/15/09 | No |
| FAQ Documenting Unhandled Exceptions | 05/21/07 | No |
| FAQ Handling ABO Incompatible Transfusion Situations | 09/23/14 | No |
| FAQ How to File a New Service Request for Changes to VBECS | RETIRED | RETIRED |
| FAQ How to take faster screen shots from Remote Desktop Session | 12/17/14 | No |
| FAQ Identifying a Comparable Blood Product Code | 09/03/21 | No |
| FAQ Local Facilities | 06/12/20 | No |
| FAQ Modification Target Not Available | 11/28/17 | No |
| FAQ Multidivisional QC Rack Workaround | 11/28/17 | No |
| FAQ Order Status Clarification | 11/15/24 | Yes |
| FAQ Polyspecific AHG Not Used for Testing | 01/29/10 | No |
| FAQ Proper Use of PIV Card with VBECS | 09/14/21 | No |
| FAQ QC Setup | 01/29/10 | No |
| FAQ Retesting QC | 11/15/24 | Yes |
| FAQ Retrieving an Expired Order | 12/07/17 | No |
| FAQ The Difference Between Transfusion Only and Full Service Facility Types | 02/08/10 | No |
| FAQ Troubleshooting Remote Desktop Issues | 12/03/24 | Yes |
| FAQ VBECS Blood Product Hierarchy | 03/28/19 | No |
| FAQ VBECS Computer Crossmatch Decision Tree | 02/08/12 | No |
| FAQ Weak D Policy | 01/11/10 | No |
| FAQ What is VBECS Read-Only | 12/04/24 | Yes |
| FAQ Workload Assignments Single vs Batch Processing | 07/01/20 | No |

# Customer Support

Contact your Local/Regional Office of Information Technology (OIT) or Laboratory Information Manager (LIM) if you encounter VistA or CPRS connection problems and for training support before contacting the Service Desk (SD).

Please ensure local/regional VistA Support contact information is available at all times.

If you experience an FDA reportable adverse event (patient death or serious injury) that VBECS may have caused or contributed to, contact the Service Desk directly to enter a ticket for Blood Bank software support.

Service Desk Contact

Call the Service Desk REDACTED

Refer to the VBECS Getting Help page on the VBECS SharePoint site for up-to-date information about contacting the Service Desk.

## References

ISBT 128 Standard Technical Specification v 6.2.2

ISBT 128 Standard Terminology for Medical Products of Human Origin, v 7.75

United States Consensus Standard for the Uniform Labeling of Blood and Blood Components Using ISBT 128, v3.0.0

Standards for Blood Banks and Transfusion Services, 34th Edition

AABB Technical Manual, 21st Edition

## VBECS SharePoint Site

The VBECS SharePoint site provides a location for additional information related to the VBECS application such as FAQs, installation status, remote desktop and report share shortcuts, and release history.

REDACTED

# Installation Qualification (IQ) Documentation

| Test Account Upgrade to VBECS 2.4.0 Revision B |  |
| --- | --- |
| Required Patch Installation | VistA Patch: None Server Patch: VBECS 2.4.0 Revision B |
| Installation Process for VBECS 2.4.0 Rev B | Installed by the BBM team, all sites will be installed in one day. The date of installation will be communicated through a ListServ announcement. |
| Expected Downtime | None |
| Installation Communication for VBECS 2.4.0 Rev B | The planned completion time for all test installs will be provided on the VBECS SharePoint site. |
| Site Responsibility | - Join VBECS-L message board on ListServ, if needed. - Communicate any changes to your site supervisor and Point of Contact (POC) to the Pathology and Laboratory Management staff (PLM). |
| Site Record of Patch Installation | Sites should take and save a screenshot of the VBECS Help, about window after the Test installation is performed. This displays the VBECS updated version information. |

| Production Account Upgrade to VBECS 2.4.0 Revision B |  |
| --- | --- |
| Required Patch Installation | VistA Patch: None Server Patch: VBECS 2.4.0 Revision B |
| Installation Process for VBECS 2.4.0 Rev B | Installed by the BBM team, all sites will be installed in one day. The date of installation will be communicated through a ListServ announcement. |
| Expected Downtime | You will be able to use VBECS during the patch installation. VBECS installation on your production account only lasts a few minutes and is invisible to users. If you encounter any disruption or errors while viewing or printing reports during the installation, try again after a few minutes. |
| Installation Communication for VBECS 2.4.0 Rev B | The status of production installs will be provided on the VBECS SharePoint site. |
| Site Responsibility | - Perform local validation, training, and set-up requirements prior to Production installation. - Join VBECS-L message board on ListServ, if needed. - Communicate any changes to your site supervisor and POC to the PLM. |
| Site Record of Patch Installation | Sites should take and save a screenshot of the VBECS Help, about window after Production installation is performed. This displays the VBECS updated version information. |

# Validation Planning

The following is a flowchart to help assess changes for validation planning.

# Table 1: Updates by Option

| #### ID | #### Option | Problem Summary | Change Summary | Validation Scenario | Change applies to my facility. (Y/N) | Local risk Assessment (Low, Med, High) | SOP revision required. If yes, identify it. | Staff training needed. (Y/N) | Scenarios or validation must be performed. (Y/N) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | CPRS Blood Bank Report 208042 | Transfusion Reaction Workup (TRW) comments over 510 are truncated. | TRW comments over 510 characters are no longer being truncated. | Record a TRW and go to the Finalize TRW page. In the Transfusion Reaction Comments section, select any Comment and enter more than 510 characters (maximum 1000 characters) in the Details section and save. View the TRW report to validate that the text entered in the Details section is no longer being truncated. Go to CPRS and view the Blood Bank Report. In the Transfusion Reactions section, validate that the text entered in the Details section is no longer being truncated. |  |  |  |  |  |
| 2 | Patient History Report 65343 | Implicated units appear multiple times when a TRW contains multiple implicated units each associated with a comment. | Implicated units appear once in the Transfusion Reactions section in the Patient History Report. | Create a TRW that contains multiple implicated units each associated with a comment. Run the Patient History Report and view the Transfusion Reactions section to validate that the implicated units no longer display multiple times. |  |  |  |  |  |
| 3 | Patient Testing Worklist Blood Availability 258406 | Patient Testing Worklist and Blood Availability reports were taking several minutes to generate for patients with multiple automated TAS tests. | Performance improvements to Patient Testing Worklist and Blood Availability reports for patients with multiple historical automated TAS tests. | This scenario is applicable only to sites with performance issues with the Patient Testing Worklist and Blood Availability reports. Prior to installation, run the Patient Testing Worklist and Blood Availability reports and record how long it takes to receive reports for a patient with a lot of historical TAS results. After installation, run the Patient Testing Worklist and Blood Availability reports and time how long it takes to receive results. |  |  |  |  |  |
| 4 | Process Outgoing Shipment / Print Invoice191657 | RBC units appear on the invoice as contaminated in certain cases when it is not contaminated | RBC units that are not contaminated no longer appear as contaminated. | Log in 2 PLATELET units that allow setting the =>2ml RBC Contamination flag (e.g., E3056V00) Activate the product code if needed. For Unit 1: check the CMV box but leave the =>2ml RBC Contamination box unchecked. Fill in any remaining required fields and add Unit 1 to the shipment. For Unit 2: check the CMV box AND the =>2ml RBC Contamination box. Fill in any remaining required fields and add Unit 2 to the shipment. Start a new Outgoing Shipment (Shipments -> Outgoing Shipment -> New Invoice). Add Unit 1 and Unit 2 to the invoice and click Save and Yes to confirm the shipment. Click Print Invoice to view the invoice and verify that only Unit2 indicates =>2ml RBC Contamination. Verify both units contain Special Testing: CMV Negative, but only Unit 2 displays =>2ml RBC Contamination. |  |  |  |  |  |
| 5 | Throughout VBECS 170263 | Update blood product table with new ICCBBA blood product codes from 7.69.0 through 7.75.4. | New blood products were added to VBECS. For a full list of new product codes see Table 2. The list of all possible non-pooled modifications for new blood products can be found on the VBECS SharePoint Site. | Validate only blood products that you may receive from your supplier. Use Tools, Blood Products to associate the blood products with an active facility. |  |  |  |  |  |
| 6 | Throughout VBECS 158622 | Update antigen type names to match 21st Edition of the AABB Technical Manual. | Minor updates to antigen type names: - V(w) to Vw - removed brackets - Jo(a ) to Jo(a) - removed extra spacing - Gya to Gy(a) - added brackets | No validation is required. |  |  |  |  |  |
| 7 | VBECS Database 180433 | Excess Address records exist with no corresponding Facility records. | Address parameter table records with no corresponding Facility parameter table records have been deleted. | No validation is required. Note: These records are not visible to the user. |  |  |  |  |  |
| 8 | VBECS SharePoint Site | SharePoint site permission changes as mandated by the OIT department | Permissions for access to the VBECS SharePoint site have been updated as mandated by the OIT department. | Validate that the appropriate team members have access to the VBECS SharePoint site. |  |  |  |  |  |
| 9 | Windows Taskbar | The VBECS application server taskbar icons are grouped, which can hide some prompts for user input, such a PIV certificate selection. | When applicable, dialog or message boxes will be seen as separate items on the taskbar of the VBECS server. | No validation is required |  |  |  |  |  |

# Table 2: New Blood Products

| Component | Product Code | Modifiable | Product Type | Long Name | Short Name | Maximum Storage Hours |
| --- | --- | --- | --- | --- | --- | --- |
| OTHER | EB078 | Yes | POOLED PLATELET-RICH BUFFY COAT | POOLED PLATELET-RICH BUFFY COAT¦CPD/XX/20-24C¦Irradiated¦ResLeu:<1E6 | BUFFY PltRch Pool CPD IRD LUKOPR | 4 |
| CRYOPRECIPITATE | EB079 | Yes | Thawed POOLED CRYOPRECIPITATE | Thawed POOLED CRYOPRECIPITATE¦NS/XX/rt¦6 units¦Methylene blue-treated | CRYO Pool Thaw NS 6U Methyl Blue | 6 |
| CRYOPRECIPITATE | EB080 | Yes | Thawed POOLED CRYOPRECIPITATE | Thawed POOLED CRYOPRECIPITATE¦NS/XX/rt¦7 units¦Methylene blue-treated | CRYO Pool Thaw NS 7U Methyl Blue | 6 |
| CRYOPRECIPITATE | EB081 | Yes | Thawed POOLED CRYOPRECIPITATE | Thawed POOLED CRYOPRECIPITATE¦NS/XX/rt¦8 units¦Methylene blue-treated | CRYO Pool Thaw NS 8U Methyl Blue | 6 |
| PLATELETS | EB082 | No | Frozen Apheresis PLATELETS | Frozen Apheresis PLATELETS¦ACD-A/XX/<=-18C¦For mnf:noninjectable¦<3E11 plts | PLT AFR Froz ACD-A MNFNI | 8760 |
| PLATELETS | EB083 | No | Frozen Apheresis PLATELETS | Frozen Apheresis PLATELETS¦ACD-A/XX/<=-18C¦For mnf:noninjectable¦1st container¦<3E11 plts | PLT AFR Froz ACD-A MNFNI 1C | 8760 |
| PLATELETS | EB084 | No | Frozen Apheresis PLATELETS | Frozen Apheresis PLATELETS¦ACD-A/XX/<=-18C¦For mnf:noninjectable¦2nd container¦<3E11 plts | PLT AFR Froz ACD-A MNFNI 2C | 8760 |
| PLATELETS | EB085 | No | Frozen Apheresis PLATELETS | Frozen Apheresis PLATELETS¦ACD-A/XX/<=-18C¦For mnf:noninjectable¦3rd container¦<3E11 plts | PLT AFR Froz ACD-A MNFNI 3C | 8760 |
| PLATELETS | EB086 | No | Frozen Apheresis PLATELETS | Frozen Apheresis PLATELETS¦ACD-A/XX/<=-18C¦For mnf:noninjectable¦Irradiated¦<3E11 plts | PLT AFR Froz ACD-A MNFNI IRD | 8760 |
| PLATELETS | EB087 | No | Frozen Apheresis PLATELETS | Frozen Apheresis PLATELETS¦ACD-A/XX/<=-18C¦For mnf:noninjectable¦Irradiated¦1st container¦<3E11 plts | PLT AFR Froz ACD-A MNFNI IRD 1C | 8760 |
| PLATELETS | EB088 | No | Frozen Apheresis PLATELETS | Frozen Apheresis PLATELETS¦ACD-A/XX/<=-18C¦For mnf:noninjectable¦Irradiated¦2nd container¦<3E11 plts | PLT AFR Froz ACD-A MNFNI IRD 2C | 8760 |
| PLATELETS | EB089 | No | Frozen Apheresis PLATELETS | Frozen Apheresis PLATELETS¦ACD-A/XX/<=-18C¦For mnf:noninjectable¦Irradiated¦3rd container¦<3E11 plts | PLT AFR Froz ACD-A MNFNI IRD 3C | 8760 |
| RED BLOOD CELLS | EB090 | No | RED BLOOD CELLS | RED BLOOD CELLS¦CP2D/XX/refg¦Not for tx or mnf¦LowVol:anticoag not adj | RBC CP2D NMT LowVol | 504 |
| RED BLOOD CELLS | EB091 | No | RED BLOOD CELLS | RED BLOOD CELLS¦CP2D/XX/refg¦Not for tx or mnf¦ResLeu:<5E6¦LowVol:anticoag not adj | RBC CP2D NMT LUKOPR LowVol | 504 |
| RED BLOOD CELLS | EB092 | No | RED BLOOD CELLS | RED BLOOD CELLS¦CP2D>AS3/XX/refg¦Not for tx or mnf¦LowVol:anticoag not adj | RBC CP2D>AS3 NMT LowVol | 1008 |
| RED BLOOD CELLS | EB093 | No | RED BLOOD CELLS | RED BLOOD CELLS¦CP2D>AS3/XX/refg¦Not for tx or mnf¦ResLeu:<5E6¦LowVol:anticoag not adj | RBC CP2D>AS3 NMT LUKOPR LowVol | 1008 |
| PLATELETS | EB094 | No | Frozen Apheresis PLATELETS | Frozen Apheresis PLATELETS¦ACD-A/XX/<=-18C¦For mnf:noninjectable¦ResLeu:<5E6¦1st container¦<3E11 plts | PLT AFR Froz ACD-A MNFNI LUKOPR 1C | 8760 |
| PLATELETS | EB095 | No | Frozen Apheresis PLATELETS | Frozen Apheresis PLATELETS¦ACD-A/XX/<=-18C¦For mnf:noninjectable¦ResLeu:<5E6¦2nd container¦<3E11 plts | PLT AFR Froz ACD-A MNFNI LUKOPR 2C | 8760 |
| PLATELETS | EB096 | No | Frozen Apheresis PLATELETS | Frozen Apheresis PLATELETS¦ACD-A/XX/<=-18C¦For mnf:noninjectable¦ResLeu:<5E6¦3rd container¦<3E11 plts | PLT AFR Froz ACD-A MNFNI LUKOPR 3C | 8760 |
| PLATELETS | EB097 | No | Frozen Apheresis PLATELETS | Frozen Apheresis PLATELETS¦ACD-A/XX/<=-18C¦For mnf:noninjectable¦Irradiated¦ResLeu:<5E6¦1st container¦<3E11 plts | PLT AFR Froz ACD-A MNFNI IRD LUKOPR 1C | 8760 |
| PLATELETS | EB098 | No | Frozen Apheresis PLATELETS | Frozen Apheresis PLATELETS¦ACD-A/XX/<=-18C¦For mnf:noninjectable¦Irradiated¦ResLeu:<5E6¦2nd container¦<3E11 plts | PLT AFR Froz ACD-A MNFNI IRD LUKOPR 2C | 8760 |
| PLATELETS | EB099 | No | Frozen Apheresis PLATELETS | Frozen Apheresis PLATELETS¦ACD-A/XX/<=-18C¦For mnf:noninjectable¦Irradiated¦ResLeu:<5E6¦3rd container¦<3E11 plts | PLT AFR Froz ACD-A MNFNI IRD LUKOPR 3C | 8760 |
| RED BLOOD CELLS | EB100 | Yes | RED BLOOD CELLS | RED BLOOD CELLS¦CPD>AS5/500mL/refg¦RBC irradiated¦ResLeu:<5E6¦Supernat rem/Plasma added | RBC CPD>AS5 RBC IRD LUKOPR SR/PA | 672 |
| RED BLOOD CELLS | EB101 | Yes | Washed RED BLOOD CELLS | Washed RED BLOOD CELLS¦None/XX/refg¦Open¦RBC irradiated¦ResLeu:<5E6¦Plasma added | RBC Wash None OPN RBC IRD LUKOPR PA | 24 |
| WHOLE BLOOD | EB102 | Yes | WHOLE BLOOD | WHOLE BLOOD¦CPD/200mL/refg | WB CPD | 504 |
| WHOLE BLOOD | EB103 | Yes | WHOLE BLOOD | WHOLE BLOOD¦CPD/400mL/refg | WB CPD | 504 |
| RED BLOOD CELLS | EB104 | Yes | RED BLOOD CELLS | RED BLOOD CELLS¦CPD>MAP/200mL/refg | RBC CPD>MAP | 1008 |
| RED BLOOD CELLS | EB105 | Yes | RED BLOOD CELLS | RED BLOOD CELLS¦CPD>MAP/200mL/refg¦ResLeu:<5E6 | RBC CPD>MAP LUKOPR | 1008 |
| RED BLOOD CELLS | EB106 | Yes | RED BLOOD CELLS | RED BLOOD CELLS¦CPD>MAP/200mL/refg¦Irradiated¦ResLeu:<5E6 | RBC CPD>MAP IRD LUKOPR | 672 |
| RED BLOOD CELLS | EB107 | Yes | RED BLOOD CELLS | RED BLOOD CELLS¦CPD>MAP/400mL/refg | RBC CPD>MAP | 1008 |
| RED BLOOD CELLS | EB108 | Yes | RED BLOOD CELLS | RED BLOOD CELLS¦CPD>MAP/400mL/refg¦ResLeu:<5E6 | RBC CPD>MAP LUKOPR | 1008 |
| RED BLOOD CELLS | EB109 | Yes | RED BLOOD CELLS | RED BLOOD CELLS¦CPD>MAP/400mL/refg¦Irradiated¦ResLeu:<5E6 | RBC CPD>MAP IRD LUKOPR | 672 |
| FRESH FROZEN PLASMA | EB110 | Yes | FRESH FROZEN PLASMA | FRESH FROZEN PLASMA¦CPD/XX/<=-20C¦ResLeu:<5E6 | FFP CPD LUKOPR | 8760 |
| FRESH FROZEN PLASMA | EB111 | Yes | Apheresis FRESH FROZEN PLASMA | Apheresis FRESH FROZEN PLASMA¦ACD-A/XX/<=-20C¦ResLeu:<5E6 | FFP AFR ACD-A LUKOPR | 8760 |
| PLATELETS | EB112 | Yes | POOLED PLATELETS | POOLED PLATELETS¦CP2D/XX/20-24C | PLT Pool CP2D | 120 |
| PLATELETS | EB113 | Yes | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A/XX/20-24C¦Open¦Irradiated¦ResLeu:<5E6¦1st container¦Approx 300 E9 plts | PLT AFR ACD-A OPN IRD LUKOPR 1C | 4 |
| RED BLOOD CELLS | EB114 | No | RED BLOOD CELLS | RED BLOOD CELLS¦CPD>SAGM/500mL/refg¦ResLeu:<5E6¦Supernat rem/SD-Plasma added | RBC CPD>SAGM LUKOPR SR/SD-PA | 1008 |
| RED BLOOD CELLS | EB115 | Yes | Washed RED BLOOD CELLS | Washed RED BLOOD CELLS¦None/XX/refg¦Open¦Irradiated¦ResLeu:<5E6¦SD-Plasma added | RBC Wash None OPN IRD LUKOPR SD-PA | 24 |
| RED BLOOD CELLS | EB116 | No | RED BLOOD CELLS | RED BLOOD CELLS¦SAGM/XX/refg¦Open¦ResLeu:<5E6¦Supernat rem/SD-Plasma added | RBC SAGM OPN LUKOPR SR/SD-PA | 24 |
| RED BLOOD CELLS | EB117 | No | RED BLOOD CELLS | RED BLOOD CELLS¦SAGM/XX/refg¦Open¦Irradiated¦ResLeu:<5E6¦Supernat rem/SD-Plasma added | RBC SAGM OPN IRD LUKOPR SR/SD-PA | 24 |
| RED BLOOD CELLS | EB118 | Yes | RED BLOOD CELLS | RED BLOOD CELLS¦CPD/XX/refg¦Irradiated¦ResLeu:<5E6¦Hct 0.75-0.9 | RBC CPD IRD LUKOPR Hct 0.75-0.9 | 504 |
| PLATELETS | EB119 | No | Washed Apheresis PLATELETS | Washed Apheresis PLATELETS¦PAS-E/XX/20-24C¦ResLeu:<5E6 | PLT AFR Wash PAS-E LUKOPR | 4 |
| PLATELETS | EB120 | No | Washed Apheresis PLATELETS | Washed Apheresis PLATELETS¦PAS-E/XX/20-24C¦Irradiated¦ResLeu:<5E6 | PLT AFR Wash PAS-E IRD LUKOPR | 4 |
| PLATELETS | EB121 | Yes | Thawed Apheresis PLATELETS | Thawed Apheresis PLATELETS¦DMSO>PAS-E/XX/20-24C¦Irradiated¦ResLeu:<5E6 | PLT AFR Thaw DMSO>PAS-E IRD LUKOPR | 4 |
| RED BLOOD CELLS | EB122 | Yes | RED BLOOD CELLS | RED BLOOD CELLS¦CP2D>AS3/500mL/refg¦Open¦RBC irradiated¦ResLeu:<5E6¦Supernat rem/Plasma added | RBC CP2D>AS3 OPN RBC IRD LUKOPR SR/PA | 24 |
| RED BLOOD CELLS | EB123 | Yes | RED BLOOD CELLS | RED BLOOD CELLS¦CPD>AS5/500mL/refg¦Open¦RBC irradiated¦ResLeu:<5E6¦Supernat rem/Plasma added | RBC CPD>AS5 OPN RBC IRD LUKOPR SR/PA | 24 |
| WHOLE BLOOD | EB124 | Yes | WHOLE BLOOD | WHOLE BLOOD¦NS/XX/rt | WB NS | 24 |
| PLATELETS | EB125 | No | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A/XX/20-24C¦Not for tx or mnf¦<3E11 plts | PLT AFR ACD-A NMT | 120 |
| PLATELETS | EB126 | No | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A/XX/20-24C¦Not for tx or mnf¦1st container¦<3E11 plts | PLT AFR ACD-A NMT 1C | 120 |
| PLATELETS | EB127 | No | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A/XX/20-24C¦Not for tx or mnf¦2nd container¦<3E11 plts | PLT AFR ACD-A NMT 2C | 120 |
| PLATELETS | EB128 | No | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A/XX/20-24C¦Not for tx or mnf¦3rd container¦<3E11 plts | PLT AFR ACD-A NMT 3C | 120 |
| PLATELETS | EB129 | No | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A/XX/20-24C¦Not for tx or mnf¦ResLeu:<5E6¦1st container¦<3E11 plts | PLT AFR ACD-A NMT LUKOPR 1C | 120 |
| PLATELETS | EB130 | No | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A/XX/20-24C¦Not for tx or mnf¦ResLeu:<5E6¦2nd container¦<3E11 plts | PLT AFR ACD-A NMT LUKOPR 2C | 120 |
| PLATELETS | EB131 | No | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A/XX/20-24C¦Not for tx or mnf¦ResLeu:<5E6¦3rd container¦<3E11 plts | PLT AFR ACD-A NMT LUKOPR 3C | 120 |
| PLATELETS | EB132 | No | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A/XX/20-24C¦Not for tx or mnf¦Psoralen-treated | PLT AFR ACD-A NMT Psoraln | 120 |
| PLATELETS | EB133 | No | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A/XX/20-24C¦Not for tx or mnf¦1st container¦Psoralen-treated | PLT AFR ACD-A NMT 1C Psoraln | 120 |
| PLATELETS | EB134 | No | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A/XX/20-24C¦Not for tx or mnf¦2nd container¦Psoralen-treated | PLT AFR ACD-A NMT 2C Psoraln | 120 |
| PLATELETS | EB135 | No | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A/XX/20-24C¦Not for tx or mnf¦3rd container¦Psoralen-treated | PLT AFR ACD-A NMT 3C Psoraln | 120 |
| PLATELETS | EB136 | Yes | POOLED PLATELETS | POOLED PLATELETS¦CPD/XX/20-24C¦Open¦ResLeu:<5E6¦4 units | PLT Pool CPD OPN LUKOPR 4U | 4 |
| PLATELETS | EB137 | Yes | POOLED PLATELETS | POOLED PLATELETS¦CPD/XX/20-24C¦ResLeu:<5E6¦Plasma reduced¦4 units | PLT Pool CPD LUKOPR PLS Red 4U | 120 |
| PLATELETS | EB138 | Yes | POOLED PLATELETS | POOLED PLATELETS¦CPD/XX/20-24C¦Open¦ResLeu:<5E6¦Plasma reduced¦4 units | PLT Pool CPD OPN LUKOPR PLS Red 4U | 4 |
| PLATELETS | EB139 | Yes | POOLED PLATELETS | POOLED PLATELETS¦CPD/XX/20-24C¦Open¦ResLeu:<5E6¦5 units | PLT Pool CPD OPN LUKOPR 5U | 4 |
| PLATELETS | EB140 | Yes | POOLED PLATELETS | POOLED PLATELETS¦CPD/XX/20-24C¦ResLeu:<5E6¦Plasma reduced¦5 units | PLT Pool CPD LUKOPR PLS Red 5U | 120 |
| PLATELETS | EB141 | Yes | POOLED PLATELETS | POOLED PLATELETS¦CPD/XX/20-24C¦Open¦ResLeu:<5E6¦Plasma reduced¦5 units | PLT Pool CPD OPN LUKOPR PLS Red 5U | 4 |
| PLATELETS | EB142 | Yes | POOLED PLATELETS | POOLED PLATELETS¦CPD/XX/20-24C¦Open¦ResLeu:<5E6¦6 units | PLT Pool CPD OPN LUKOPR 6U | 4 |
| PLATELETS | EB143 | Yes | POOLED PLATELETS | POOLED PLATELETS¦CPD/XX/20-24C¦ResLeu:<5E6¦Plasma reduced¦6 units | PLT Pool CPD LUKOPR PLS Red 6U | 120 |
| PLATELETS | EB144 | Yes | POOLED PLATELETS | POOLED PLATELETS¦CPD/XX/20-24C¦Open¦ResLeu:<5E6¦Plasma reduced¦6 units | PLT Pool CPD OPN LUKOPR PLS Red 6U | 4 |
| PLATELETS | EB145 | Yes | POOLED PLATELETS | POOLED PLATELETS¦CPD/XX/20-24C¦Open¦Irradiated¦ResLeu:<5E6¦4 units | PLT Pool CPD OPN IRD LUKOPR 4U | 4 |
| PLATELETS | EB146 | Yes | POOLED PLATELETS | POOLED PLATELETS¦CPD/XX/20-24C¦Irradiated¦ResLeu:<5E6¦Plasma reduced¦4 units | PLT Pool CPD IRD LUKOPR PLS Red 4U | 120 |
| PLATELETS | EB147 | Yes | POOLED PLATELETS | POOLED PLATELETS¦CPD/XX/20-24C¦Open¦Irradiated¦ResLeu:<5E6¦Plasma reduced¦4 units | PLT Pool CPD OPN IRD LUKOPR PLS Red 4U | 4 |
| PLATELETS | EB148 | Yes | POOLED PLATELETS | POOLED PLATELETS¦CPD/XX/20-24C¦Open¦Irradiated¦ResLeu:<5E6¦5 units | PLT Pool CPD OPN IRD LUKOPR 5U | 4 |
| PLATELETS | EB149 | Yes | POOLED PLATELETS | POOLED PLATELETS¦CPD/XX/20-24C¦Irradiated¦ResLeu:<5E6¦Plasma reduced¦5 units | PLT Pool CPD IRD LUKOPR PLS Red 5U | 120 |
| PLATELETS | EB150 | Yes | POOLED PLATELETS | POOLED PLATELETS¦CPD/XX/20-24C¦Open¦Irradiated¦ResLeu:<5E6¦6 units | PLT Pool CPD OPN IRD LUKOPR 6U | 4 |
| PLATELETS | EB151 | Yes | POOLED PLATELETS | POOLED PLATELETS¦CPD/XX/20-24C¦Irradiated¦ResLeu:<5E6¦Plasma reduced¦6 units | PLT Pool CPD IRD LUKOPR PLS Red 6U | 120 |
| PLATELETS | EB152 | Yes | POOLED PLATELETS | POOLED PLATELETS¦CPD/XX/20-24C¦Open¦ResLeu:<5E6¦<3E11 plts | PLT Pool CPD OPN LUKOPR | 4 |
| PLATELETS | EB153 | Yes | POOLED PLATELETS | POOLED PLATELETS¦CPD/XX/20-24C¦ResLeu:<5E6¦Plasma reduced¦<3E11 plts | PLT Pool CPD LUKOPR PLS Red | 120 |
| PLATELETS | EB154 | Yes | POOLED PLATELETS | POOLED PLATELETS¦CPD/XX/20-24C¦Open¦ResLeu:<5E6¦Plasma reduced¦<3E11 plts | PLT Pool CPD OPN LUKOPR PLS Red | 4 |
| PLATELETS | EB155 | Yes | POOLED PLATELETS | POOLED PLATELETS¦CPD/XX/20-24C¦Open¦Irradiated¦ResLeu:<5E6¦<3E11 plts | PLT Pool CPD OPN IRD LUKOPR | 4 |
| PLATELETS | EB156 | Yes | POOLED PLATELETS | POOLED PLATELETS¦CPD/XX/20-24C¦Irradiated¦ResLeu:<5E6¦Plasma reduced¦<3E11 plts | PLT Pool CPD IRD LUKOPR PLS Red | 120 |
| PLATELETS | EB157 | Yes | POOLED PLATELETS | POOLED PLATELETS¦CPD/XX/20-24C¦Open¦Irradiated¦ResLeu:<5E6¦Plasma reduced¦<3E11 plts | PLT Pool CPD OPN IRD LUKOPR PLS Red | 4 |
| RED BLOOD CELLS | EB158 | Yes | Washed Apheresis RED BLOOD CELLS | Washed Apheresis RED BLOOD CELLS¦None/XX/refg¦Open¦RBC irradiated¦ResLeu:<5E6¦Plasma added | RBC AFR Wash None OPN RBC IRD LUKOPR PA | 24 |
| RED BLOOD CELLS | EB159 | Yes | Washed Apheresis RED BLOOD CELLS | Washed Apheresis RED BLOOD CELLS¦None/XX/refg¦Open¦RBC irradiated¦ResLeu:<5E6¦Plasma added¦1st container | RBC AFR Wash None OPN RBC IRD LUKOPR PA 1C | 24 |
| RED BLOOD CELLS | EB160 | Yes | Washed Apheresis RED BLOOD CELLS | Washed Apheresis RED BLOOD CELLS¦None/XX/refg¦Open¦RBC irradiated¦ResLeu:<5E6¦Plasma added¦2nd container | RBC AFR Wash None OPN RBC IRD LUKOPR PA 2C | 24 |
| RED BLOOD CELLS | EB161 | Yes | RED BLOOD CELLS | RED BLOOD CELLS¦CP2D>AS3/500mL/refg¦ResLeu:<5E6¦O2/CO2 reduced | RBC CP2D>AS3 LUKOPR O2/CO2 Red | 1008 |
| RED BLOOD CELLS | EB162 | Yes | RED BLOOD CELLS | RED BLOOD CELLS¦CP2D>AS3/500mL/refg¦Irradiated¦ResLeu:<5E6¦O2/CO2 reduced | RBC CP2D>AS3 IRD LUKOPR O2/CO2 Red | 672 |
| RED BLOOD CELLS | EB163 | Yes | Deglycerolized RED BLOOD CELLS | Deglycerolized RED BLOOD CELLS¦NS/XX/refg¦Open¦ResLeu:<5E6¦Supernat rem¦Hct 0.7-0.8 | RBC Deg NS OPN LUKOPR SR Hct 0.7-0.8 | 24 |
| PLATELETS | EB164 | Yes | POOLED PLATELETS | POOLED PLATELETS¦CPD>PAS-E/XX/20-24C¦ResLeu:<5E6¦Supernat reduced¦Buffy coat plts prep¦Psoralen-treated | PLT Pool CPD>PAS-E LUKOPR SUPRed BUFCT Psoraln | 120 |
| PLATELETS | EB165 | No | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A>PAS-E/XX/20-24C¦ResLeu:<5E6¦Supernat reduced¦Psoralen-treated | PLT AFR ACD-A>PAS-E LUKOPR SUPRed Psoraln | 120 |
| PLATELETS | EB166 | No | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A>PAS-E/XX/20-24C¦ResLeu:<5E6¦Supernat reduced¦1st container¦Psoralen-treated | PLT AFR ACD-A>PAS-E LUKOPR SUPRed 1C Psoraln | 120 |
| PLATELETS | EB167 | No | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A>PAS-E/XX/20-24C¦ResLeu:<5E6¦Supernat reduced¦2nd container¦Psoralen-treated | PLT AFR ACD-A>PAS-E LUKOPR SUPRed 2C Psoraln | 120 |
| PLATELETS | EB168 | No | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A>PAS-E/XX/20-24C¦ResLeu:<5E6¦Supernat reduced¦3rd container¦Psoralen-treated | PLT AFR ACD-A>PAS-E LUKOPR SUPRed 3C Psoraln | 120 |
| PLATELETS | EB169 | Yes | POOLED PLATELETS | POOLED PLATELETS¦CPD>PAS-E/XX/20-24C¦ResLeu:<5E6¦Supernat rem/Albumin added¦Buffy coat plts prep¦Psoralen-treated | PLT Pool CPD>PAS-E LUKOPR SR/ALB Add BUFCT Psoraln | 120 |
| PLATELETS | EB170 | No | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A>PAS-E/XX/20-24C¦ResLeu:<5E6¦Supernat rem/Albumin added¦Psoralen-treated | PLT AFR ACD-A>PAS-E LUKOPR SR/ALB Add Psoraln | 120 |
| PLATELETS | EB171 | No | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A>PAS-E/XX/20-24C¦ResLeu:<5E6¦Supernat rem/Albumin added¦1st container¦Psoralen-treated | PLT AFR ACD-A>PAS-E LUKOPR SR/ALB Add 1C Psoraln | 120 |
| PLATELETS | EB172 | No | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A>PAS-E/XX/20-24C¦ResLeu:<5E6¦Supernat rem/Albumin added¦2nd container¦Psoralen-treated | PLT AFR ACD-A>PAS-E LUKOPR SR/ALB Add 2C Psoraln | 120 |
| PLATELETS | EB173 | No | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A>PAS-E/XX/20-24C¦ResLeu:<5E6¦Supernat rem/Albumin added¦3rd container¦Psoralen-treated | PLT AFR ACD-A>PAS-E LUKOPR SR/ALB Add 3C Psoraln | 120 |
| PLATELETS | EB174 | Yes | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A/XX/20-24C¦Open¦ResLeu:<5E6¦Plasma reduced/Saline added¦<3E11 plts | PLT AFR ACD-A OPN LUKOPR PLS Red/SAL Add | 4 |
| PLATELETS | EB175 | Yes | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A/XX/20-24C¦Open¦ResLeu:<5E6¦Plasma reduced/Saline added¦1st container¦<3E11 plts | PLT AFR ACD-A OPN LUKOPR PLS Red/SAL Add 1C | 4 |
| PLATELETS | EB176 | Yes | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A/XX/20-24C¦Open¦ResLeu:<5E6¦Plasma reduced/Saline added¦2nd container¦<3E11 plts | PLT AFR ACD-A OPN LUKOPR PLS Red/SAL Add 2C | 4 |
| PLATELETS | EB177 | Yes | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A/XX/20-24C¦Open¦ResLeu:<5E6¦Plasma reduced/Saline added¦3rd container¦<3E11 plts | PLT AFR ACD-A OPN LUKOPR PLS Red/SAL Add 3C | 4 |
| PLATELETS | EB178 | Yes | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A/XX/20-24C¦Open¦Irradiated¦ResLeu:<5E6¦Plasma reduced/Saline added¦<3E11 plts | PLT AFR ACD-A OPN IRD LUKOPR PLS Red/SAL Add | 4 |
| PLATELETS | EB179 | Yes | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A/XX/20-24C¦Open¦Irradiated¦ResLeu:<5E6¦Plasma reduced/Saline added¦1st container¦<3E11 plts | PLT AFR ACD-A OPN IRD LUKOPR PLS Red/SAL Add 1C | 4 |
| PLATELETS | EB180 | Yes | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A/XX/20-24C¦Open¦Irradiated¦ResLeu:<5E6¦Plasma reduced/Saline added¦3rd container¦<3E11 plts | PLT AFR ACD-A OPN IRD LUKOPR PLS Red/SAL Add 3C | 4 |
| PLATELETS | EB181 | Yes | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A/XX/20-24C¦Open¦ResLeu:<5E6¦Plasma reduced¦3rd container¦<3E11 plts¦Psoralen-treated | PLT AFR ACD-A OPN LUKOPR PLS Red 3C Psoraln | 4 |
| PLATELETS | EB182 | Yes | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A/XX/20-24C¦Open¦ResLeu:<5E6¦Plasma reduced/Saline added¦1st container¦<3E11 plts¦Psoralen-treated | PLT AFR ACD-A OPN LUKOPR PLS Red/SAL Add 1C Psoraln | 4 |
| PLATELETS | EB183 | Yes | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A/XX/20-24C¦Open¦ResLeu:<5E6¦Plasma reduced/Saline added¦2nd container¦<3E11 plts¦Psoralen-treated | PLT AFR ACD-A OPN LUKOPR PLS Red/SAL Add 2C Psoraln | 4 |
| PLATELETS | EB184 | Yes | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A/XX/20-24C¦Open¦ResLeu:<5E6¦Plasma reduced/Saline added¦3rd container¦<3E11 plts¦Psoralen-treated | PLT AFR ACD-A OPN LUKOPR PLS Red/SAL Add 3C Psoraln | 4 |
| PLATELETS | EB185 | Yes | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A/XX/20-24C¦Irradiated¦ResLeu:<5E6¦Plasma reduced¦4th container¦<3E11 plts | PLT AFR ACD-A IRD LUKOPR PLS Red 4C | 120 |
| PLATELETS | EB186 | Yes | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A/XX/20-24C¦Open¦Irradiated¦ResLeu:<5E6¦Plasma reduced/Saline added¦4th container¦<3E11 plts | PLT AFR ACD-A OPN IRD LUKOPR PLS Red/SAL Add 4C | 4 |
| WHOLE BLOOD | EB187 | Yes | WHOLE BLOOD | WHOLE BLOOD¦CPD/450mL/rt¦ResLeu:<1E6 | WB CPD LUKOPR | 504 |
| WHOLE BLOOD | EB188 | Yes | WHOLE BLOOD | WHOLE BLOOD¦CPD/450mL/rt¦Irradiated¦ResLeu:<1E6 | WB CPD IRD LUKOPR | 504 |
| PLATELETS | EB189 | No | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A>PAS-E/XX/20-24C¦ResLeu:<1E6¦1st container¦Psoralen-treated | PLT AFR ACD-A>PAS-E LUKOPR 1C Psoraln | 120 |
| PLATELETS | EB190 | No | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A>PAS-E/XX/20-24C¦ResLeu:<1E6¦2nd container¦Psoralen-treated | PLT AFR ACD-A>PAS-E LUKOPR 2C Psoraln | 120 |
| PLATELETS | EB191 | No | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A>PAS-E/XX/20-24C¦ResLeu:<1E6¦Supernat reduced¦1st container¦Psoralen-treated | PLT AFR ACD-A>PAS-E LUKOPR SUPRed 1C Psoraln | 120 |
| PLATELETS | EB192 | No | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A>PAS-E/XX/20-24C¦ResLeu:<1E6¦Supernat reduced¦2nd container¦Psoralen-treated | PLT AFR ACD-A>PAS-E LUKOPR SUPRed 2C Psoraln | 120 |
| PLATELETS | EB193 | No | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A>PAS-E/XX/20-24C¦ResLeu:<1E6¦1st container¦Approx 120 E9 plts¦Psoralen-treated | PLT AFR ACD-A>PAS-E LUKOPR 1C Psoraln | 120 |
| PLATELETS | EB194 | No | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A>PAS-E/XX/20-24C¦ResLeu:<1E6¦2nd container¦Approx 120 E9 plts¦Psoralen-treated | PLT AFR ACD-A>PAS-E LUKOPR 2C Psoraln | 120 |
| FRESH FROZEN PLASMA | EB195 | Yes | Apheresis FRESH FROZEN PLASMA | Apheresis FRESH FROZEN PLASMA¦ACD-A/XX/<=-25C¦ResLeu:<1E6¦>=200mL<400mL¦Frozen <=6h¦1st container | FFP AFR ACD-A LUKOPR 1C | 8760 |
| FRESH FROZEN PLASMA | EB196 | Yes | Apheresis FRESH FROZEN PLASMA | Apheresis FRESH FROZEN PLASMA¦ACD-A/XX/<=-25C¦ResLeu:<1E6¦>=200mL<400mL¦Frozen <=6h¦2nd container | FFP AFR ACD-A LUKOPR 2C | 8760 |
| FRESH FROZEN PLASMA | EB197 | Yes | Apheresis FRESH FROZEN PLASMA | Apheresis FRESH FROZEN PLASMA¦ACD-A/XX/<=-25C¦ResLeu:<1E6¦>=200mL<400mL¦Frozen <=6h¦3rd container | FFP AFR ACD-A LUKOPR 3C | 8760 |
| FRESH FROZEN PLASMA | EB198 | Yes | Apheresis FRESH FROZEN PLASMA | Apheresis FRESH FROZEN PLASMA¦ACD-A/XX/<=-25C¦ResLeu:<1E6¦<125 mL¦Frozen <=6h | FFP AFR ACD-A LUKOPR | 8760 |
| FRESH FROZEN PLASMA | EB199 | No | Apheresis FRESH FROZEN PLASMA | Apheresis FRESH FROZEN PLASMA¦ACD-A/XX/<=-25C¦ResLeu:<1E6¦<125 mL¦Frozen <=6h¦Quar:>=4m/retested | FFP AFR ACD-A LUKOPR QUAR | 8760 |
| FRESH FROZEN PLASMA | EB200 | No | Thawed Apheresis FRESH FROZEN PLASMA | Thawed Apheresis FRESH FROZEN PLASMA¦ACD-A/XX/rt¦ResLeu:<1E6¦<125 mL¦Frozen <=6h¦Quar:>=4m/retested | FFP AFR Thaw ACD-A LUKOPR QUAR | 24 |
| FRESH FROZEN PLASMA | EB201 | No | Apheresis FRESH FROZEN PLASMA | Apheresis FRESH FROZEN PLASMA¦ACD-A/XX/<=-25C¦ResLeu:<1E6¦>=200mL<400mL¦Frozen <=6h¦1st container¦Quar:>=4m/retested | FFP AFR ACD-A LUKOPR 1C QUAR | 8760 |
| FRESH FROZEN PLASMA | EB202 | No | Apheresis FRESH FROZEN PLASMA | Apheresis FRESH FROZEN PLASMA¦ACD-A/XX/<=-25C¦ResLeu:<1E6¦>=200mL<400mL¦Frozen <=6h¦2nd container¦Quar:>=4m/retested | FFP AFR ACD-A LUKOPR 2C QUAR | 8760 |
| FRESH FROZEN PLASMA | EB203 | No | Apheresis FRESH FROZEN PLASMA | Apheresis FRESH FROZEN PLASMA¦ACD-A/XX/<=-25C¦ResLeu:<1E6¦>=200mL<400mL¦Frozen <=6h¦3rd container¦Quar:>=4m/retested | FFP AFR ACD-A LUKOPR 3C QUAR | 8760 |
| FRESH FROZEN PLASMA | EB204 | No | Thawed Apheresis FRESH FROZEN PLASMA | Thawed Apheresis FRESH FROZEN PLASMA¦ACD-A/XX/rt¦ResLeu:<1E6¦>=200mL<400mL¦Frozen <=6h¦1st container¦Quar:>=4m/retested | FFP AFR Thaw ACD-A LUKOPR 1C QUAR | 24 |
| FRESH FROZEN PLASMA | EB205 | No | Thawed Apheresis FRESH FROZEN PLASMA | Thawed Apheresis FRESH FROZEN PLASMA¦ACD-A/XX/rt¦ResLeu:<1E6¦>=200mL<400mL¦Frozen <=6h¦2nd container¦Quar:>=4m/retested | FFP AFR Thaw ACD-A LUKOPR 2C QUAR | 24 |
| FRESH FROZEN PLASMA | EB206 | No | Thawed Apheresis FRESH FROZEN PLASMA | Thawed Apheresis FRESH FROZEN PLASMA¦ACD-A/XX/rt¦ResLeu:<1E6¦>=200mL<400mL¦Frozen <=6h¦3rd container¦Quar:>=4m/retested | FFP AFR Thaw ACD-A LUKOPR 3C QUAR | 24 |
| PLATELETS | EB207 | Yes | POOLED PLATELETS | POOLED PLATELETS¦CPD>PAS-E/XX/20-24C¦ResLeu:<1E6¦Buffy coat plts prep¦Psoralen-treated | PLT Pool CPD>PAS-E LUKOPR BUFCT Psoraln | 120 |
| PLATELETS | EB208 | Yes | POOLED PLATELETS | POOLED PLATELETS¦CPD>PAS-E/XX/20-24C¦ResLeu:<1E6¦Supernat reduced¦Buffy coat plts prep¦Psoralen-treated | PLT Pool CPD>PAS-E LUKOPR SUPRed BUFCT Psoraln | 120 |
| PLATELETS | EB209 | No | Frozen POOLED PLATELETS | Frozen POOLED PLATELETS¦CPD/XX/<=-18C¦For mnf:noninjectable¦ResLeu:<5E6¦4 units | PLT Pool Froz CPD MNFNI LUKOPR 4U | 8760 |
| PLATELETS | EB210 | No | Frozen POOLED PLATELETS | Frozen POOLED PLATELETS¦CPD/XX/<=-18C¦For mnf:noninjectable¦ResLeu:<5E6¦5 units | PLT Pool Froz CPD MNFNI LUKOPR 5U | 8760 |
| PLATELETS | EB211 | No | Frozen POOLED PLATELETS | Frozen POOLED PLATELETS¦CPD/XX/<=-18C¦For mnf:noninjectable¦ResLeu:<5E6¦6 units | PLT Pool Froz CPD MNFNI LUKOPR 6U | 8760 |
| PLATELETS | EB212 | No | Frozen POOLED PLATELETS | Frozen POOLED PLATELETS¦CPD/XX/<=-18C¦For mnf:noninjectable¦Irradiated¦ResLeu:<5E6¦4 units | PLT Pool Froz CPD MNFNI IRD LUKOPR 4U | 8760 |
| PLATELETS | EB213 | No | Frozen POOLED PLATELETS | Frozen POOLED PLATELETS¦CPD/XX/<=-18C¦For mnf:noninjectable¦Irradiated¦ResLeu:<5E6¦5 units | PLT Pool Froz CPD MNFNI IRD LUKOPR 5U | 8760 |
| PLATELETS | EB214 | No | Frozen POOLED PLATELETS | Frozen POOLED PLATELETS¦CPD/XX/<=-18C¦For mnf:noninjectable¦Irradiated¦ResLeu:<5E6¦6 units | PLT Pool Froz CPD MNFNI IRD LUKOPR 6U | 8760 |
| RED BLOOD CELLS | EB215 | Yes | RED BLOOD CELLS | RED BLOOD CELLS¦CP2D>AS3/500mL/refg¦RBC irradiated¦ResLeu:NS¦Supernat rem/Plasma added | RBC CP2D>AS3 RBC IRD LUKOPR SR/PA | 672 |
| RED BLOOD CELLS | EB216 | Yes | Apheresis RED BLOOD CELLS | Apheresis RED BLOOD CELLS¦CP2D>AS3/XX/refg¦RBC irradiated¦ResLeu:NS¦Supernat rem/Plasma added¦1st container | RBC AFR CP2D>AS3 RBC IRD LUKOPR SR/PA 1C | 672 |
| RED BLOOD CELLS | EB217 | Yes | Apheresis RED BLOOD CELLS | Apheresis RED BLOOD CELLS¦CP2D>AS3/XX/refg¦RBC irradiated¦ResLeu:NS¦Supernat rem/Plasma added¦2nd container | RBC AFR CP2D>AS3 RBC IRD LUKOPR SR/PA 2C | 672 |
| RED BLOOD CELLS | EB218 | Yes | Apheresis RED BLOOD CELLS | Apheresis RED BLOOD CELLS¦ACD-A>AS1/XX/refg¦RBC irradiated¦ResLeu:NS¦Supernat rem/Plasma added¦1st container | RBC AFR ACD-A>AS1 RBC IRD LUKOPR SR/PA 1C | 672 |
| RED BLOOD CELLS | EB219 | Yes | Apheresis RED BLOOD CELLS | Apheresis RED BLOOD CELLS¦ACD-A>AS1/XX/refg¦RBC irradiated¦ResLeu:NS¦Supernat rem/Plasma added¦2nd container | RBC AFR ACD-A>AS1 RBC IRD LUKOPR SR/PA 2C | 672 |
| PLATELETS | EB220 | No | POOLED PLATELETS | POOLED PLATELETS¦PAS-E/XX/20-24C¦ResLeu:<1E6¦Supernat rem/Plasma added¦Buffy coat plts prep¦4 units | PLT Pool PAS-E LUKOPR SR/PA BUFCT 4U | 120 |
| PLATELETS | EB221 | No | POOLED PLATELETS | POOLED PLATELETS¦PAS-E/XX/20-24C¦ResLeu:<1E6¦Supernat rem/Plasma added¦Buffy coat plts prep¦5 units | PLT Pool PAS-E LUKOPR SR/PA BUFCT 5U | 120 |
| PLATELETS | EB222 | No | POOLED PLATELETS | POOLED PLATELETS¦PAS-E/XX/20-24C¦ResLeu:<1E6¦Supernat rem/Plasma added¦Buffy coat plts prep¦6 units | PLT Pool PAS-E LUKOPR SR/PA BUFCT 6U | 120 |
| PLATELETS | EB223 | Yes | POOLED PLATELETS | POOLED PLATELETS¦CPD/XX/20-24C¦ResLeu:<5E6¦<3E11 plts¦Bacterial monitoring >=36h | PLT Pool CPD LUKOPR BacMtr >=36h | 120 |
| PLATELETS | EB224 | Yes | POOLED PLATELETS | POOLED PLATELETS¦CPD/XX/20-24C¦Irradiated¦ResLeu:<5E6¦<3E11 plts¦Bacterial monitoring >=36h | PLT Pool CPD IRD LUKOPR BacMtr >=36h | 120 |
| WHOLE BLOOD | EB225 | Yes | WHOLE BLOOD | WHOLE BLOOD¦CPD/XX/refg¦ResLeu:<5E6¦Platelets spared | WB CPD LUKOPR PLT S | 504 |
| FRESH FROZEN PLASMA | EB226 | Yes | Apheresis FRESH FROZEN PLASMA | Apheresis FRESH FROZEN PLASMA¦NaCitrate/XX/<=-25C¦ResLeu:<1E6¦3rd container | FFP AFR NaCit LUKOPR 3C | 8760 |
| FRESH FROZEN PLASMA | EB227 | Yes | Apheresis FRESH FROZEN PLASMA | Apheresis FRESH FROZEN PLASMA¦NaCitrate/XX/<=-25C¦Irradiated¦ResLeu:<1E6¦3rd container | FFP AFR NaCit IRD LUKOPR 3C | 8760 |
| FRESH FROZEN PLASMA | EB228 | Yes | Thawed Apheresis FRESH FROZEN PLASMA | Thawed Apheresis FRESH FROZEN PLASMA¦NaCitrate/XX/refg¦Irradiated¦ResLeu:<1E6¦3rd container | FFP AFR Thaw NaCit IRD LUKOPR 3C | 24 |
| FRESH FROZEN PLASMA | EB229 | Yes | Apheresis PLASMA | Apheresis PLASMA¦ACD-A/XX/<=-18C¦1st container | PLASMA AFR ACD-A 1C | 8760 |
| FRESH FROZEN PLASMA | EB230 | Yes | Apheresis PLASMA | Apheresis PLASMA¦ACD-A/XX/<=-18C¦2nd container | PLASMA AFR ACD-A 2C | 8760 |
| FRESH FROZEN PLASMA | EB231 | Yes | Apheresis PLASMA | Apheresis PLASMA¦ACD-A/XX/<=-18C¦3rd container | PLASMA AFR ACD-A 3C | 8760 |
| FRESH FROZEN PLASMA | EB232 | Yes | Apheresis PLASMA | Apheresis PLASMA¦ACD-A/XX/<=-18C¦4th container | PLASMA AFR ACD-A 4C | 8760 |
| FRESH FROZEN PLASMA | EB233 | Yes | Liquid PLASMA | Liquid PLASMA¦CP2D/XX/refg¦ResLeu:<5E6 | PLASMA Liq CP2D LUKOPR | 624 |
| RED BLOOD CELLS | EB234 | Yes | Washed RED BLOOD CELLS | Washed RED BLOOD CELLS¦None/XX/refg¦Open¦ResLeu:<5E6¦Supernat reduced | RBC Wash None OPN LUKOPR SUPRed | 24 |
| RED BLOOD CELLS | EB235 | Yes | Washed RED BLOOD CELLS | Washed RED BLOOD CELLS¦None/XX/refg¦Open¦Irradiated¦ResLeu:<5E6¦Supernat reduced | RBC Wash None OPN IRD LUKOPR SUPRed | 24 |
| RED BLOOD CELLS | EB236 | Yes | Washed Apheresis RED BLOOD CELLS | Washed Apheresis RED BLOOD CELLS¦None/XX/refg¦Open¦ResLeu:<5E6¦Supernat reduced | RBC AFR Wash None OPN LUKOPR SUPRed | 24 |
| RED BLOOD CELLS | EB237 | Yes | Washed Apheresis RED BLOOD CELLS | Washed Apheresis RED BLOOD CELLS¦None/XX/refg¦Open¦ResLeu:<5E6¦Supernat reduced¦1st container | RBC AFR Wash None OPN LUKOPR SUPRed 1C | 24 |
| RED BLOOD CELLS | EB238 | Yes | Washed Apheresis RED BLOOD CELLS | Washed Apheresis RED BLOOD CELLS¦None/XX/refg¦Open¦ResLeu:<5E6¦Supernat reduced¦2nd container | RBC AFR Wash None OPN LUKOPR SUPRed 2C | 24 |
| RED BLOOD CELLS | EB239 | Yes | Washed Apheresis RED BLOOD CELLS | Washed Apheresis RED BLOOD CELLS¦None/XX/refg¦Open¦Irradiated¦ResLeu:<5E6¦Supernat reduced | RBC AFR Wash None OPN IRD LUKOPR SUPRed | 24 |
| RED BLOOD CELLS | EB240 | Yes | Washed Apheresis RED BLOOD CELLS | Washed Apheresis RED BLOOD CELLS¦None/XX/refg¦Open¦Irradiated¦ResLeu:<5E6¦Supernat reduced¦1st container | RBC AFR Wash None OPN IRD LUKOPR SUPRed 1C | 24 |
| RED BLOOD CELLS | EB241 | Yes | Washed Apheresis RED BLOOD CELLS | Washed Apheresis RED BLOOD CELLS¦None/XX/refg¦Open¦Irradiated¦ResLeu:<5E6¦Supernat reduced¦2nd container | RBC AFR Wash None OPN IRD LUKOPR SUPRed 2C | 24 |
| PLATELETS | EB242 | Yes | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A>PAS-C/XX/refg¦Open¦ResLeu:<5E6¦Psoralen-treated | PLT AFR ACD-A>PAS-C OPN LUKOPR Psoraln | 4 |
| PLATELETS | EB243 | Yes | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A>PAS-C/XX/refg¦Open¦ResLeu:<5E6¦1st container¦Psoralen-treated | PLT AFR ACD-A>PAS-C OPN LUKOPR 1C Psoraln | 4 |
| PLATELETS | EB244 | Yes | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A>PAS-C/XX/refg¦Open¦ResLeu:<5E6¦2nd container¦Psoralen-treated | PLT AFR ACD-A>PAS-C OPN LUKOPR 2C Psoraln | 4 |
| PLATELETS | EB245 | Yes | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A>PAS-C/XX/refg¦Open¦ResLeu:<5E6¦3rd container¦Psoralen-treated | PLT AFR ACD-A>PAS-C OPN LUKOPR 3C Psoraln | 4 |
| OTHER | EB246 | Yes | PLATELET-RICH BUFFY COAT | PLATELET-RICH BUFFY COAT¦CPDA-1/500mL/20-24C | BUFFY PltRch CPDA-1 | 24 |
| OTHER | EB247 | Yes | PLATELET-RICH BUFFY COAT | PLATELET-RICH BUFFY COAT¦CPDA-1/500mL/20-24C¦Irradiated | BUFFY PltRch CPDA-1 IRD | 24 |
| PLATELETS | EB248 | Yes | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A>PAS-B/XX/refg¦ResLeu:<5E6¦Psoralen-treated | PLT AFR ACD-A>PAS-B LUKOPR Psoraln | 120 |
| PLATELETS | EB249 | Yes | POOLED PLATELETS | POOLED PLATELETS¦CPD>PAS-E/XX/refg¦ResLeu:<5E6¦Buffy coat plts prep¦Psoralen-treated | PLT Pool CPD>PAS-E LUKOPR BUFCT Psoraln | 120 |
| PLATELETS | EB250 | Yes | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A>PAS-E/XX/refg¦ResLeu:<5E6¦1st container | PLT AFR ACD-A>PAS-E LUKOPR 1C | 120 |
| PLATELETS | EB251 | Yes | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A>PAS-E/XX/refg¦ResLeu:<5E6¦2nd container | PLT AFR ACD-A>PAS-E LUKOPR 2C | 120 |
| PLATELETS | EB252 | Yes | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A>PAS-E/XX/refg¦Irradiated¦ResLeu:<5E6¦1st container | PLT AFR ACD-A>PAS-E IRD LUKOPR 1C | 120 |
| PLATELETS | EB253 | Yes | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A>PAS-E/XX/refg¦Irradiated¦ResLeu:<5E6¦2nd container | PLT AFR ACD-A>PAS-E IRD LUKOPR 2C | 120 |
| PLATELETS | EB254 | Yes | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A>PAS-E/XX/refg¦ResLeu:<5E6¦1st container¦Psoralen-treated | PLT AFR ACD-A>PAS-E LUKOPR 1C Psoraln | 120 |
| PLATELETS | EB255 | Yes | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A>PAS-E/XX/refg¦ResLeu:<5E6¦2nd container¦Psoralen-treated | PLT AFR ACD-A>PAS-E LUKOPR 2C Psoraln | 120 |
| FRESH FROZEN PLASMA | EB256 | Yes | Apheresis PLASMA | Apheresis PLASMA¦ACD-A/XX/<=-18C¦Frozen <=24h¦1st container¦Psoralen-treated | PLASMA AFR ACD-A 1C Psoraln | 8760 |
| FRESH FROZEN PLASMA | EB257 | Yes | Apheresis PLASMA | Apheresis PLASMA¦ACD-A/XX/<=-18C¦Frozen <=24h¦2nd container¦Psoralen-treated | PLASMA AFR ACD-A 2C Psoraln | 8760 |
| FRESH FROZEN PLASMA | EB258 | Yes | Apheresis PLASMA | Apheresis PLASMA¦ACD-A/XX/<=-18C¦Frozen <=24h¦3rd container¦Psoralen-treated | PLASMA AFR ACD-A 3C Psoraln | 8760 |
| PLATELETS | EB259 | Yes | POOLED PLATELETS | POOLED PLATELETS¦CPD>PAS-E/XX/20-24C¦ResLeu:<1E6¦Supernat reduced | PLT Pool CPD>PAS-E LUKOPR SUPRed | 120 |
| PLATELETS | EB260 | Yes | POOLED PLATELETS | POOLED PLATELETS¦CPD>PAS-E/XX/20-24C¦Irradiated¦ResLeu:<1E6¦Supernat reduced | PLT Pool CPD>PAS-E IRD LUKOPR SUPRed | 120 |
| PLATELETS | EB261 | No | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A>PAS-E/XX/20-24C¦Open¦ResLeu:<5E6¦Plasma reduced/Saline added¦1st container¦Psoralen-treated | PLT AFR ACD-A>PAS-E OPN LUKOPR PLS Red/SAL Add 1C Psoraln | 4 |
| PLATELETS | EB262 | No | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A>PAS-E/XX/20-24C¦Open¦ResLeu:<5E6¦Plasma reduced/Saline added¦2nd container¦Psoralen-treated | PLT AFR ACD-A>PAS-E OPN LUKOPR PLS Red/SAL Add 2C Psoraln | 4 |
| OTHER | EB263 | Yes | PLATELET-RICH BUFFY COAT | PLATELET-RICH BUFFY COAT¦CPD/350mL/20-24C | BUFFY PltRch CPD | 24 |
| OTHER | EB264 | Yes | PLATELET-RICH BUFFY COAT | PLATELET-RICH BUFFY COAT¦CPD/350mL/20-24C¦Irradiated | BUFFY PltRch CPD IRD | 24 |
| PLATELETS | EB265 | No | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A>PAS-E/XX/20-24C¦Open¦ResLeu:<5E6¦Plasma reduced¦1st container | PLT AFR ACD-A>PAS-E OPN LUKOPR PLS Red 1C | 4 |
| PLATELETS | EB266 | No | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A>PAS-E/XX/20-24C¦Open¦ResLeu:<5E6¦Plasma reduced¦2nd container | PLT AFR ACD-A>PAS-E OPN LUKOPR PLS Red 2C | 4 |
| PLATELETS | EB267 | No | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A>PAS-E/XX/20-24C¦Open¦ResLeu:<5E6¦Plasma reduced¦1st container¦Psoralen-treated | PLT AFR ACD-A>PAS-E OPN LUKOPR PLS Red 1C Psoraln | 4 |
| PLATELETS | EB268 | No | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A>PAS-E/XX/20-24C¦Open¦ResLeu:<5E6¦Plasma reduced¦2nd container¦Psoralen-treated | PLT AFR ACD-A>PAS-E OPN LUKOPR PLS Red 2C Psoraln | 4 |
| PLATELETS | EB269 | No | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A>PAS-E/XX/20-24C¦Open¦Irradiated¦ResLeu:<5E6¦Plasma reduced¦1st container | PLT AFR ACD-A>PAS-E OPN IRD LUKOPR PLS Red 1C | 4 |
| PLATELETS | EB270 | No | Apheresis PLATELETS | Apheresis PLATELETS¦ACD-A>PAS-E/XX/20-24C¦Open¦Irradiated¦ResLeu:<5E6¦Plasma reduced¦2nd container | PLT AFR ACD-A>PAS-E OPN IRD LUKOPR PLS Red 2C | 4 |
| FRESH FROZEN PLASMA | EB271 | Yes | Thawed PLASMA | Thawed PLASMA¦CPD/XX/refg¦1st container | PLASMA Thaw CPD 1C | 120 |
| FRESH FROZEN PLASMA | EB272 | Yes | Thawed PLASMA | Thawed PLASMA¦CPD/XX/refg¦2nd container | PLASMA Thaw CPD 2C | 120 |
| FRESH FROZEN PLASMA | EB273 | Yes | Thawed PLASMA | Thawed PLASMA¦CPD/XX/refg¦3rd container | PLASMA Thaw CPD 3C | 120 |
| PLATELETS | EB274 | Yes | POOLED PLATELETS | POOLED PLATELETS¦CPD>PAS-E/XX/20-24C¦Open¦ResLeu:<5E6¦Plasma reduced/Saline added¦Buffy coat plts prep¦Psoralen-treated | PLT Pool CPD>PAS-E OPN LUKOPR PLS Red/SAL Add BUFCT Psoraln | 4 |
| RED BLOOD CELLS | EB275 | Yes | RED BLOOD CELLS | RED BLOOD CELLS¦CPD>AS5/450mL/refg¦ResLeu:NS | RBC CPD>AS5 LUKOPR | 1008 |
| RED BLOOD CELLS | EB276 | Yes | RED BLOOD CELLS | RED BLOOD CELLS¦CPD>AS5/XX/refg¦ResLeu:NS¦LowVol:anticoag not adj | RBC CPD>AS5 LUKOPR LowVol | 1008 |
| RED BLOOD CELLS | EB277 | Yes | RED BLOOD CELLS | RED BLOOD CELLS¦CPD>AS5/XX/refg¦ResLeu:<1E6¦LowVol:anticoag not adj | RBC CPD>AS5 LUKOPR LowVol | 1008 |
| FRESH FROZEN PLASMA | EB278 | Yes | FRESH FROZEN PLASMA | FRESH FROZEN PLASMA¦CPD/450mL/<=-18C¦Frozen <=8h | FFP CPD | 8760 |
| FRESH FROZEN PLASMA | EB279 | Yes | Apheresis FRESH FROZEN PLASMA | Apheresis FRESH FROZEN PLASMA¦NaCitrate/XX/<=-25C¦Frozen <=18h¦Psoralen-treated | FFP AFR NaCit Psoraln | 8760 |
| FRESH FROZEN PLASMA | EB280 | Yes | POOLED FRESH FROZEN PLASMA | POOLED FRESH FROZEN PLASMA¦CPD/XX/<=-18C¦From 5 donors | FFP Pool CPD 5DNR | 8760 |
| RED BLOOD CELLS | EB281 | Yes | RED BLOOD CELLS | RED BLOOD CELLS¦CPD>AS1/500mL/refg¦Open¦RBC irradiated¦ResLeu:NS¦Plasma added | RBC CPD>AS1 OPN RBC IRD LUKOPR PA | 24 |
| RED BLOOD CELLS | EB282 | No | RED BLOOD CELLS | RED BLOOD CELLS¦SAGM/XX/refg¦Irradiated¦ResLeu:<1E6¦LowVol:anticoag adj | RBC SAGM IRD LUKOPR LowVol | 672 |
| FRESH FROZEN PLASMA | X0014 | Yes | Thawed SOLVENT DETERGENT POOLED PLASMA | Thawed SOLVENT DETERGENT POOLED PLASMA¦NS/NS/refg¦ABO independent | PLASMA Pool Thaw Det Sol NS ABO ind | 24 |

This is the last page of VBECS 2.4.0 Rev B Release Notes.