MAG\*3.0\*185 Deployment, Installation, Back-out, and Rollback Plan



**Version 19.0 MAG\*3.0\*185**

## Department of Veterans Affairs

**Office of Information and Technology (OI&T)**

**MAG\*3.0\*185 Deployment, Installation, Back-out, and Rollback Plan**

**Property of the US Government**

This is a controlled document. No changes to this document may be made without the express written consent of the VistA Imaging development group.

While every effort has been made to assure the accuracy of the information provided, this document may include technical inaccuracies and/or typographical errors. Changes are periodically made to the information herein and incorporated into new editions of this document.

Product names mentioned in this document may be trademarks or registered trademarks of their respective companies, and are hereby acknowledged.

VistA Imaging Product Development Department of Veterans Affairs Internet: REDACTED

VA intranet: REDACTED

**Revision History**

|  |  |  |
| --- | --- | --- |
| **Date** | **Rev** | **Notes** |
| Oct 14 2011 | 1 | Created for MAG\*3.0\*104. WPR complete Aug 31; labeled for release Oct 14. REDACTED |
| May 07 2012 | 2 | Updated for MAG\*3.0\*122. REDACTED |
| July 25 2013 | 3 | Updated for MAG\*3.0\*034/116/118, MAG\*3.0\*119, MAG\*3.0\*127, MAG\*3.0\*129. REDACTED |
| September 09  2013 | 4 | Updated for MAG\*3.0\*130, REDACTED |
| May 10 2016 | 5 | Updated for MAG\*3.0\*162, REDACTED |
| February 14  2017 | 6 | Updated for MAG 3.0\*170, REDACTED |
| March 24  2017 | 7 | Date updates REDACTED |
| April 27 2017 | 8 | Updated port information REDACTED |
| July 05 2017 | 9 | MAG\*3.0\*170 Installer Updates REDACTED |
| July 26 2017 | 10 | MAG\*3.0\*170 Updates REDACTED |
| August 01  2017 | 11 | MAG\*3.0\*177 Updates REDACTED |
| October 03  2017 | 12 | MAG\*3.0\*177 Updates REDACTED |
| October 25  2017 | 13 | MAG\*3.0\*185 Updates. REDACTED |

Contents

|  |  |  |
| --- | --- | --- |
| **Date** | **Rev** | **Notes** |
| November 14  2017 | 14 | Incorporated comments from reviewers regarding MAG\*3.0\*185. REDACTED |
| December 5  2017 | 15 | Updated MAG\*3.0\*185 build numbers. REDACTED |
| January 11  2018 | 16 | Additional MAG\*3.0\*185 updates. REDACTED |
| January 25  2018 | 17 | Additional MAG\*3.0\*185 Updates. REDACTED |
| February 7  2018 | 18 | Additional MAG\*3.0\*185 Updates. REDACTED |
| February 8  2018 | 19 | Additional MAG\*3.0\*185 Updates. REDACTED |

# Contents

[Introduction 1](#_bookmark0)

[Intended Audience 1](#_bookmark1)

[Terms of Use 1](#_bookmark2)

[Document Conventions 1](#_bookmark3)

[Related Information 2](#_bookmark4)

[Installing a New VIX 3](#_bookmark5)

[Preparing for a New VIX Installation 3](#_bookmark6)

[Setting up VistA 3](#_bookmark7)

[Selecting and Validating the VIX Server 3](#_bookmark8)

[Getting VIX Component Licenses 7](#_bookmark9)

[Getting a VIX Security Certificate 8](#_bookmark10)

[Java Version 8](#_bookmark11)

[New VIX Installation – Standalone Server 8](#_bookmark12)

[Preparing Passwords, Activating Components, and Staging Files 8](#_bookmark13)

[Verifying Installation Is Complete 18](#_bookmark14)

[Activating a New VIX 21](#_bookmark15)

[Updating an Existing VIX 22](#_bookmark16)

[Preparing for a VIX Update 22](#_bookmark17)

[VistA Software Dependencies 24](#_bookmark18)

[Scheduling Downtime and Impact of a VIX Update 24](#_bookmark19)

[Java Version 25](#_bookmark20)

[Performing a VIX Update – Standalone Server 25](#_bookmark21)

[Post-installation 34](#_bookmark22)

[McAfee Exclusions 34](#_bookmark23)

[Verifying VIX Operations 35](#_bookmark24)

[Verifying Access to the VIX Transaction Log 35](#_bookmark25)

[Spot-checking VIX Image Delivery 35](#_bookmark26)

[Using the VIX Installation Wizard to Reconfigure the VIX 38](#_bookmark27)

[Reconfiguring a VIX – Standalone Server 38](#_bookmark28)

[Troubleshooting 40](#_bookmark29)

[Resuming an Interrupted VIX Installation 40](#_bookmark30)

[Resuming Installation (single server VIX) 40](#_bookmark31)

[VIX Support 40](#_bookmark32)

[Back Out/Uninstall 42](#_bookmark33)

[Back Out/Uninstall Scenarios 42](#_bookmark34)

[Uninstall/Restore as part of Troubleshooting 42](#_bookmark35)

[Relocating a VIX onto a New Server 42](#_bookmark36)

[Decommissioning a VIX 43](#_bookmark37)

[Uninstalling the VIX 43](#_bookmark38)

[Stopping the VIX service 44](#_bookmark39)

[Remove VIX-related applications, accounts, directories, and variables. 44](#_bookmark40)

Contents

[Appendix A: VIX Install Checklist 48](#_bookmark41)

[VIX Install Checklist 48](#_bookmark42)

# Introduction

This document explains how to install the VistA Imaging Exchange (VIX) service. Please review the install checklist in Appendix A prior to starting the install and reference throughout.

## Intended Audience

This document is intended for VA staff responsible for managing a local VIX.

This document presumes a working knowledge of the VistA environment, VistA Imaging components and workflow, and Windows administration.

## Terms of Use

The VIX is a component of VistA Imaging and is regulated as a medical device by the Food and Drug Administration (FDA). Use of the VIX is subject to the following provisions:

|  |
| --- |
| The FDA classifies VistA Imaging, and the VIX (as a component of VistA Imaging) as a medical device. Unauthorized modifications to VistA Imaging, including the VIX, such as the installation of unapproved software, will adulterate the medical device. The use of an adulterated medical device violates US federal law (21CFR820). |
| Because software distribution/inventory management tools can install  inappropriate or unapproved software without a local administrator’s knowledge, sites must exclude the VIX server from such systems. |

## Document Conventions

This document uses the following conventions:

* Controls, options, and button names are shown in **Bold**.
* A vertical bar is used to separate successive menu choices. For example: “Click

**File | Open**” means: “Click the **File** menu; then click the **Open** option.”

* Keyboard key names are shown in bold and in brackets.
* Sample output is shown in monospace.
* Important or required information is shown in a **Note**.
* Critical information is indicated by:

Introduction

## Related Information

In addition to this manual, the following document contains information about the VIX:

REDACTED

# Installing a New VIX

This section explains how to implement a new VIX. The installation checklist in Appendix A, VIX Update Checklist, summarizes the process.

**Tip:** If you are updating an existing VIX, see the [*Updating an Existing VIX*](#_bookmark16) section.

## Preparing for a New VIX Installation

Preparing to install a new VIX involves:

* Acquiring Installation Files
  + MAG3\_0P185\_VIX\_Setup.msi and SQLExpress\_x64-12.0.2000.8.zip

should be copied to a temporary folder on the desktop.

* Setting up VistA
* Selecting and validating the server where the VIX will be installed
* Getting VIX component licenses
* Getting a VIX security certificate Specifics are covered in the following sections.

**Setting up VistA**

You must install the compatible KIDS package on the VistA system before installing the VIX server software. For information about how to install the KIDS package, see the patch description of the patch you are installing.

* If you are implementing a VIX for the first time, the MAG VIX ADMIN key, introduced in Patch MAG\*3.0\*83, must be assigned to the VistA accounts of administrators who need access to the VIX transaction log.
* While it is not required, it is recommended that sites run the MagDexter and MagKat utilities provided in patch MAG\*3.0\*98. Doing so populates DICOM series information for radiology exams acquired before the release of MAG\*3.0\*50. See the [*VistA Imaging Storage Utilities Manual*](https://www.va.gov/vdl/application.asp?appid=105) for details.

**Selecting and Validating the VIX Server**

The server where the VIX is installed must meet the FDA, hardware, and operating system/environmental requirements specified in the following sections.


**FDA Requirements**

The VIX is a component of VistA Imaging and is therefore regulated as a medical device by the Food and Drug Administration (FDA). Use of the VIX is subject to the terms of use listed in the introduction.


**Hardware Requirements**

Minimum VIX hardware requirements:

* Installation must be done on a 64bit OS machine.
* Minimum
  + 2 CPUs and 4 gigabytes of RAM
* Preferred
  + 4 CPUs and 8 gigabytes of RAM
* A dedicated local drive for the VIX cache
* Minimum
  + 200 gigabytes of disk space
* Preferred
  + 500 gigabytes of disk space
* A 1 gigabit Ethernet connection will need to be available for use by the VIX.



*Figure 1:* ***MINIMUM*** *VIX Settings*



**Port Requirements**

*Figure 2:* ***PREFERRED*** *VIX Settings*

On the server where the VIX is installed, verify that ports 343, 443, 8080, and 8443 are accessible to VA wide area network IP addresses (10.x.x.x).

**Operating System and Environment Requirements**

The VIX can run on:

* Minimum: Windows Server 2008 R2
* Preferred: Windows Server 2012 R2

**.NET Framework Requirement**

* If you are installing on Windows Server 2008 R2
  + Verify that .Net 3.5 is installed and enabled. It must be enabled for the VIX install.
  + .Net 4.5 framework might not be installed on your server. If so, this condition causes the installer to fail. To discover if you have the feature installed it should be present in *“Control Panel\Programs\Programs and Features”* as *“Microsoft .NET Framework 4.5.x”*.
  + To remedy download and install the .Net 4.5 framework. **(for convenience, a copy is placed on the FTP site)**
* If you are installing on Windows Server 2012 R2
  + Verify that .Net 3.5 is enabled. It must be enabled for the VIX install.
  + .Net 4.5 framework should be installed on your server. If it isn’t installed, the installer will fail. To discover if you have the feature installed it should be present in *“Control Panel\Programs\Programs and Features”* as

*“Microsoft .NET Framework 4.5.x”*.

* + To remedy download and install the .Net 4.5 framework. **(for convenience, a copy is placed on the FTP site)**

**Getting VIX Component Licenses**

Two toolkits used by the VIX require third-party licenses.

**Laurel Bridge DCF Toolkit**

The Laurel Bridge DICOM Connectivity Framework (DCF) toolkit is a third-party toolkit used by the VIX to convert images to and from DICOM format.

The VA has purchased an enterprise-wide license for this toolkit. To request site-specific serial numbers for this license, do the following:

1. From the server where the VIX is installed, attempt to access the Laurel Bridge activation code server at REDACTED. (A login page will display if you can access the site.)

If you cannot this access this site, enter a ticket and indicate that your site ACL (access control list) needs to be modified to access REDACTED.

1. Contact the [REDACTED](mailto:VHAVILBLicenses@va.gov) mail group. They will provide a request form and instructions for completing the form.

The serial number information provided will ultimately be used to generate an activation code (this is done during the VIX installation process). The activation code allows the Laurel Bridge DCF toolkit to be installed and used.

The Laurel Bridge activation code is linked to the hardware of the server where the toolkit is installed. If you have to replace the licensed server, email the contacts listed previously to arrange to get the serial number(s) transferred as well.

**Aware JPEG2000 Toolkit License**

The Aware JPEG2000 toolkit is a third-party toolkit used for DICOM-to-JPEG2000 image conversion by the VIX.

Sites that implement the VIX must purchase a one-time permanent license for the latest version of the Aware JPEG2000 toolkit from Aware for each server where the VIX is used

**IMPORTANT:** Do not install the Aware JPEG2000 toolkit before you run the VIX installer. The VIX installer installs the version of the Aware JPEG2000 toolkit with which the VIX has been tested. If you install the Aware JPEG2000 toolkit before running the VIX installer, this will interfere with the installation process.

**Getting a VIX Security Certificate**

Each server that hosts the VIX must be issued a security certificate. To get security certificates, send an e-mail to REDACTED and include the following information:

* Fully Qualified Domain Name (FQDN) of the VIX server name
* Contact information for the primary and backup administrators of the VIX.

Requests will typically be processed within five business days. After the request has been processed, the security certificate (one per server) will be securely transmitted to the requesting site POC for install on the site VIX.

Each security certificate received will need to be copied to the local server where the VIX software is installed. The VIX installation wizard will prompt you for this certificate during the installation process.

**Java Version**

Before beginning the install, validate the specific patch description document for the current version of Java. If the server has a different version of Java than the required version specified in patch description, please uninstall the application.

## New VIX Installation – Standalone Server


Use the following steps to install a VIX on a standalone server for the first time. Installation should take less than 30 minutes assuming that all preparation steps in previous sections are complete.



**Note** Only perform the steps in this section if you are installing the VIX on a single server.

**Preparing Passwords, Activating Components, and Staging Files**

Before starting the VIX installation process on a standalone server, do the following:

1. Verify that .NET 3.5 features and .NET 4.5 features are enabled. If Windows needs to download files, accept.





1. Prepare a password to be used for the Apache Tomcat administrator account that will be created as part of the VIX installation process.
   * This account will be rarely used; it only needs to be secured properly.
   * The password is case-sensitive and only alphanumeric characters are allowed.
2. Prepare a password for the Windows account that will be created as part of the VIX installation process.
   * This Windows account, which will be named “apachetomcat” when it is created by the VIX installer, is used to run the VIX in the Tomcat environment. This account is limited to only the functions it needs to run the VIX.
   * The password is case sensitive, must contain at least eight characters, and must contain at least one capital letter and one number.
3. Set up a service account in VistA for ROI periodic processing with the MAG DICOM VISA secondary menu option and the OR CPRS GUI CHART secondary menu option. The service account may be the same service account as the one the DICOM Gateway and the HDIG use. The credentials are required for the VIX to process ROI disclosure requests periodically, in the background and to purge completed requests.
4. Determine the email address or addresses for notification about invalid ROI periodic processing credentials. The VIX will send an email notification to the email address or addresses, if the ROI periodic processing credentials are invalid or have expired. You must specify this address when you install the VIX. You can set up a new email account for this purpose or use an existing one.
5. Copy the VIX security certificate to the primary drive (usually the C drive) of the server where the VIX will be installed.

**Standalone Server Installation**

VIX installation involves running two processes back-to-back. The first short process installs the VIX Installation Wizard. The second, longer process involves using the VIX Installation Wizard to install the actual VIX.


**Note:** These steps presume that you have already obtained a serial number for the Laurel Bridge DCF toolkit. These steps also presume the VIX can access the Laurel Bridge license server via the internet.

**To install the VIX on a standalone server**

1. Use an administrator-level account to log on to the standalone server where the VIX will be installed.
2. Copy the latest VIX installer (MAG3\_0P<number>\_VIX\_Setup and SQLExpress\_x64-12.0.2000.8.zip) to a temporary folder on the desktop.
3. Do the following to prepare the VIX Installation Wizard:
   1. Double-click the VIX set-up MAG3\_0P<number>\_VIX\_Setup.msi file.
4. When the Welcome page displays, click **Next**.
5. When the Confirm Installation page displays, click **Next**.
6. When the Installation Complete screen displays, click **Close**.

### Choose Start | All Programs | VistA Imaging Programs | VIX Installation Service Wizard.

1. Run selected program as administrator.



1. When the Welcome page for the VIX installation wizard displays, click **Next**.
2. In the **Site Number** field of the Specify the VA site... page, enter the STATION NUMBER (field (#99) in the INSTITUTION file (#4)) of your site. Then, click **Lookup Server Addresses**.
3. Verify that the site-related information retrieved by the lookup is correct. Then, click

### Next.

**Note**: The VIX server hostname will be blank and the port number will be 0; these values are populated automatically once the VIX is registered with the VistA site service.

1. When the Install the VIX Prerequisites page displays, verify that the icons for the first two items on the page are Checkmark icon .

* **<account> has Administrator role –** This line will indicate if an administrative account is being used. If not, cancel the installation and restart it using a Windows administrator account.
* **Current operating system –** This line will indicate if the proper operating system is present. If a non-supported operating system is identified, the installation cannot continue.

1. On the same page, check the line that indicates the state of the Java Runtime environment. If Install needed icon is shown, do the following:
   1. Click **Install**.
   2. Wait until the status icon the Java Runtime Environment changes to Checkmark icon . (This install of Java will take longer than previous versions and may take several minutes to complete)
2. On the same page, check the line that indicates the state of the Apache Tomcat installation. If Install needed icon is shown, do the following:
   1. Click **Install**.
   2. In the dialog box that displays, enter and confirm the Apache Tomcat password that you prepared as described in [Preparing Passwords, Activating Components, and Staging Files](#_bookmark13) previously. Then, click **OK**.
   3. Wait until the status icon for Apache Tomcat changes to Checkmark icon . (This install of Tomcat will take longer than previous versions and may take several minutes to complete)
3. On the same page, check the line that indicates the state of the Laurel Bridge toolkit. If Install needed icon is shown, do the following:
   1. Click the **Install** button next to the Laurel Bridge item. After a brief delay, the Activate DCF License window will open.

**Tip:** The Network Activation tab will be selected automatically, and about half of the boxes in the window will be pre-populated.

* 1. Enter all of the following information in the Network Activation tab:



**Note:** The Activate button will be disabled if any of the following boxes are left blank.

* + - **Product Serial Number** – the new Laurel Bridge DCF serial number (include dashes).
    - **Site** – the name of your site.
    - **CPUs** – the number of CPUs on the server hosting the VIX.
    - **Contact name** and **Contact email** – the administrator of your local VistA Imaging system.
  1. Near the bottom of the window, click **Activate**. After a brief delay, the **Status**

box will display a green “Success” message.


* 1. Click **Exit with success**. The Activate DCF License window will close and the updated Laurel Bridge toolkit will be installed (installation will only take a second or two).

1. On the same page, check the line that indicates the state of the service account. If Install needed icon is shown, do the following:
   1. Click **Create**.
   2. In the dialog box that displays, enter the Windows service account password that you prepared [Preparing Passwords, Activating Components, and Staging Files](#_bookmark13) previously. Then, click **OK**.
   3. Wait until the status icon for the service account changes to Checkmark icon .
2. On the same page, check the line that indicates the state of VIX security certificate. If Install needed icon is shown, do the following:
   1. Click **Install**.
   2. In the dialog box that displays, click **Select**.
   3. Specify the location of the security certificate setup file received from the national VistA Imaging team. Then, click **OK.**
3. Check the line that indicates VIX/Viewer Render Services. If Install needed icon is shown, do the following.
   1. Click **Install**
   2. Click **Configure Viewer/Render**



* 1. Verify the following settings:

1. Verify the viewer port is set to 343.
2. Verify Site Service host name is set to localhost.
3. Verify Site Service port is set to 8080.
4. Verify VIX Service host name is set to localhost.
5. Verify instance name is .\ SQLEXPRESS
   1. Edit image cache directory drive to the dedicated VIX cache drive. (For example "E:\VIXRenderCache")
   2. Click the **Save Configuration** button in the top right corner.
   3. Click **OK**
   4. If prompted to install the SQL server, click **OK** and select the .zip file in the temporary folder on the desktop. (Depending on your system, this step may take up to twenty minutes)




1. In the Install the VIX Prerequisites page, confirm that all the icons are Checkmark icon . Then, click **Next**.
2. In the Specify the location page, select the drive where you want the VIX configuration files to reside. Then, click **Create**.

* In most cases, you should use the same drive for the VIX configuration files and for the VIX cache.

1. In the same page, select the drive where the VIX cache will be located. Then, click


### Create.

1. Click **Next**.
2. Select the “C:\” drive for VixConfig folder and the “E:\” drive for VixCache folder. Click each **Create** button to create the folders.



VIX Service Installation Wizard 30.185.8.6060

1. Click **Next**.
2. In the Specify the Release of Information (ROI) Configuration, do the following:
   * Specify the access and verify codes for the account with the ROI periodic processing credentials. The VIX uses this account for periodic processing of ROI disclosure requests. The account must be valid VistA credentials with the MAG DICOM VISA secondary menu option and the OR CPRS GUI CHART secondary menu option. The credentials can be the credentials of the same service account that the DICOM Gateway and the HDIG use.
   * Specify the email address that gets notifications for invalid ROI periodic processing credentials. The VIX sends an email notification to the address or addresses specified in this field if the ROI periodic processing credentials are expired or invalid. You can enter several addresses, separated by a comma.




VIX Service Installation Wizard 30.180.8.6060

1. Click **Validate**.

The VIX installation program checks the information. If it detects an error, it displays a tooltip with information about the error. When it validates the configuration, **Next** becomes available.


1. In the Configure Local DoD connection page, do one of the following:
   * If your site has no local network connection to a DoD facility, click **Next** (this will be the case at most VA sites)
   * If your site has a local network connection to a DoD facility, enter connection

information for the DoD’s PACS Integrator server. After entering the connection information, click **Validate** to test the connection. Then, click **Next.**




VIX Service Installation Wizard 30.185.8.6060

1. On the Install the VIX page, click **Install**. (The information in this page is saved in C:\Program Files (x86)\VistA\Imaging\ViX Installer for future reference or troubleshooting.) This will start the installation process. It will also start the Tomcat and Viewer/Render services.



VIX Service Installation Wizard 30.185.8.6060

1. When installation is complete, click **Finish** to exit the wizard.




VIX Service Installation Wizard 30.185.8.6060

1. The VIX will start automatically, but cannot be used until it is activated and registered with the VistA Site Service. See the next section for details.


## Verifying Installation Is Complete


After you have run the VIX Installation Wizard, verify that your installation is complete. Follow these steps:

1. Go to the VIX homepage: http://<FQDN of VIX server>:8080/
2. The current version is listed in the format XX.XXX.X.X. The first two digits represent Version 3.0 of the VistA Imaging system and do not change. The next three digits are the number of the latest patch that has affected the VIX. This patch number should match the number of the VIX component you have most recently installed.



Welcome to the VISA Service, your current version is 30.185.8.6060

1. Verify VIX viewer and VIX render services are running.
2. Verify the listener tool is running.



## Activating a New VIX

After the VIX is installed, it will be inactive until it is registered with the VistA Site Service. Clinical Display workstations and VistARad workstations use the site service to determine if local and remote VIX servers are available.

After the VIX is registered in the site service, the VIX will begin to be actively used, both by clinicians at your site as well as by remote VA sites for access to locally stored images.

**Note:** Do not register the VIX with the site service until after it is installed. Registering the VIX before it is installed will cause errors and timeout issues for Clinical Display users.

### To register the VIX

1. Gather the following information:

**Primary contact name, phone, and email:**

**Backup contact name, phone, and email:**

**Site name:**

**STATION NUMBER (field #99) from INSTITUTION file (#4):**

**Network Name defined for Imaging Resources group**:

1. Enter a CA ticket to Clin 3 for site service update.
   * Paste the lines in the preceding step into the ticket.
   * Include “Add VIX server to Site Service database”
2. You will be notified, typically within five business days, when the site service registration is complete.
3. See [*Verifying VIX Operations*](#_bookmark24)

# Updating an Existing VIX

This chapter explains how to update an existing VIX server. An installation checklist that summarizes this process is on page 52.

## Preparing for a VIX Update

Preparing for a VIX update involves:

* + Acquiring Installation Files

o MAG3\_0P185\_VIX\_Setup.msi and SQLExpress\_x64-12.0.2000.8.zip


should be copied to a temporary folder on the desktop.

* + Making sure that the KIDS version installed on the VistA system matches the version the VIX requires.
  + Scheduling downtime and notifying users of the impact of a VIX update.
  + **Tip:** Aware licenses (used for DICOM-to-JPEG2000 conversion on the VIX) established for older VIXes will automatically carry over to new VIXes.
  + **Tip:** VIX-specific account passwords and security key assignments established for older VIXes will automatically carry over to new VIXes.


**Hardware Requirements**

Minimum VIX hardware requirements:

* Installation must be done on a 64bit OS machine.
* Minimum:
  + 2 CPUs and 4 gigabytes of RAM
* Preferred:
  + 4 CPUs and 8 gigabytes of RAM
* A dedicated local drive for the VIX cache
  + Minimum
    - 200 gigabytes of disk space
  + Preferred
    - 500 gigabytes of disk space
* A 1 gigabit Ethernet connection will need to be available for use by the VIX.



*Figure 3:* ***MINIMUM*** *VIX Settings*



*Figure 4:* ***PREFERRED*** *VIX Settings*

**VistA Software Dependencies**

The VIX server software requires that a compatible Imaging KIDS package be installed on VistA. For details about the compatible KIDS package and installing it, see the patch description of the specific patch.

**Scheduling Downtime and Impact of a VIX Update**

You will need to schedule downtime with appropriate personnel for the duration of the VIX installation.

**Note:** If a VIX is installed on a standalone (dedicated) server, the DICOM image acquisition is not impacted and can continue.

While the VIX server is being updated, VIX assisted functions will not be available. The following table summarizes how a VIX outage will affect clinicians:

|  |  |
| --- | --- |
| **Clinical Group** | **Impact** |
| Local Clinical Display users | DoD images will not be accessible for the duration of the VIX shutdown.  Remote VA images may be temporarily inaccessible. Clinical Display will attempt to revert to pre-VIX remote image views, but users may have to disconnect from and reconnect to remote sites, or in some cases, restart Clinical Display.  Clinicians may notice longer retrieval times for remote images for the duration of the VIX shutdown.  After the VIX is restarted, restart Clinical Display to make sure that Clinical Display is no longer using pre-VIX remote image views for remote sites. |
| Local VistARad users | Remote exam data and monitored exam lists will not be available. For additional information, refer to the documentation for VistARad. |
| Remote VA or DoD clinicians requesting your site’s images | Remote clinicians may experience transitory application errors if the VIX is shut down while it is in the middle of processing a request; the clinician may have to repeat the request.  Remote VA clinicians issuing new requests may notice longer retrieval times for the duration of the VIX shutdown.  Remote DoD clinicians will not be able to retrieve locally stored images for the duration of the VIX shutdown. |
| Local DICOM Importer users | The DICOM importer client will be unable to log into VistA and process imports for the duration that the VIX is down. |
| VIX image viewer/JLV users | Clinical users of JLV will not be able to access images or artifacts using the VIX image viewer. |

**Java Version**

Before beginning the install, validate the specific patch description document for the current version of Java. If the server has a different version of Java than the required version specified in patch description, please uninstall the application.

## Performing a VIX Update – Standalone Server

Use the following steps to update a VIX on a standalone server.

**Note:** These steps presume that you have already obtained a new Laurel Bridge DCF toolkit serial number. These steps also presume the VIX can access the Laurel Bridge license server via the internet.


1. Use an administrator-level account to log on to the standalone server where the VIX will be installed.
2. Copy the VIX installation files (MAG3\_0P<number>\_VIX\_Setup.msi and

SQLExpress\_x64-12.0.2000.8.zip) to a local folder on the server.


1. Do the following to prepare the VIX Installation Wizard:
   1. Double-click the VIX installation file.
   2. When the Welcome page displays, click **Next**.
   3. When the Confirm Installation page displays, click **Next**.
   4. When the Installation Complete screen displays, click **Close**.

### Choose Start | All Programs | VistA Imaging Programs | VIX Installation Service Wizard.

1. Run selected program as administrator.



1. When the Welcome page displays, click **Next**. Then, when you are prompted to do so, uninstall the pre-existing VIX software. (The wizard will gracefully stop the VIX service before performing the uninstall.)
2. When the Uninstall complete message displays near the top of the page, click **Next**.



**Update VIX components from 30.177.5.6044 to 30.185.8.6060**

Click uninstall button to remove existing VIX web applications.


VIX Service Installation Wizard 30.185.8.6060

*Figure 5: Updating from 177 to 185*

1. In the **Site Number** field of the Specify the VA Site... page, verify that the **Site Number** box shows your STATION NUMBER (field (#99) in the INSTITUTION file (#4)).
2. Confirm the connection by clicking **Lookup Server Address**. Then click **Next**.
3. In the VIX Prerequisites page:
4. Check the line that indicates the state of the Java Runtime environment. If Install needed icon is shown, do the following:
   * Click **Install**.
   * Wait until the status icon the Java Runtime Environment changes to Checkmark icon . (This install of Java will take longer than previous versions and may take several minutes to complete)


* Check the line that indicates the state of the Apache Tomcat installation. If Install needed icon is shown, do the following:
  + Click **Install**.
  + In the dialog box that displays, enter and confirm the Apache Tomcat password that you prepared as described in [Preparing Passwords, Activating Components, and Staging Files](#_bookmark13) previously. Then, click **OK**.
  + Wait until the status icon for Apache Tomcat changes to Checkmark icon . (This install of Tomcat will take longer than previous versions and may take several minutes to complete)

1. On the same page, check the line that indicates the state of the service account. If Install needed icon is shown, do the following:
   * Click **Assign**.



VIX Service Installation Wizard 30.185.8.6060



* + If your service user did not change, click **Yes.** If you need to reassign the password, click **No.**




VIX Service Installation Wizard 30.185.8.6060

* + If you answered **No,** in the dialog box that displays, enter the Windows service account password that you prepared [Preparing Passwords, Activating Components, and Staging Files](#_bookmark13) previously. Then, click **OK**.
  + Wait until the status icon for the service account changes to Checkmark icon .


1. Then click **Next** three more times to accept defaults in the following three pages.
2. Check the line that indicates VIX/Viewer Render Services.
   1. If Install needed icon is shown, click **Install.** If Checkmark icon is shown, click **Reinstall.**
   2. Click **Configure Viewer/Render**



* 1. Verify the following settings:
     1. Verify the viewer port is set to 343.
     2. Verify Site Service host name is set to localhost.
     3. Verify Site Service port is set to 8080.
     4. Verify VIX Service host name is set to localhost.
     5. Verify instance name is .\ SQLEXPRESS
  2. Edit image cache directory drive to the dedicated VIX cache drive. (For example "E:\VIXRenderCache")
  3. Click the **Save Configuration** button in the top right corner.
  4. Click **OK**
  5. If prompted to install the SQL server, click **OK** and select the .zip file in the temporary folder on the desktop. (Depending on your system, this step may take up to twenty minutes)




1. Click **Next.**
2. Select the “C:\” drive for VixConfig folder and the dedicated VIX cache drive for VixCache folder. (For example, “E:\VIXCache”) Click each **Create** button to create the folders.



VIX Service Installation Wizard 30.185.8.6060

1. In the Specify the Release of Information (ROI) Configuration, do the following:

* Specify the access and verify codes for the account with the ROI periodic processing credentials. The VIX uses this account for periodic processing of ROI disclosure requests. The account must be valid VistA credentials with the MAG DICOM VISA secondary menu option and the OR CPRS GUI CHART secondary menu option. The credentials can be the credentials of the same service account that the DICOM Gateway and the HDIG use.
* Specify the email address that gets notifications for invalid ROI periodic processing credentials. The VIX sends an email notification to the address or addresses specified in this field if the ROI periodic processing credentials are expired or invalid. You can enter several addresses, separated by a comma.



VIX Service Installation Wizard 30.185.8.6060


1. Click **Validate**.

The VIX installation program checks the information. If it detects an error, it displays a tooltip with information about the error. When it validates the configuration, **Next** becomes available.

1. In the Configure Local DoD connection page, do one of the following:


* If your site has no local network connection to a DoD facility, click **Next** (this will be the case at most VA sites)
* If your site has a local network connection to a DoD facility, enter connection

information for the DoD’s PACS Integrator server. After entering the connection information, click **Validate** to test the connection. Then, click **Next.**



VIX Service Installation Wizard 30.185.8.6060


1. On the Install the VIX page, click **Install**. (The information in this page is saved in C:\Program Filesx (x86)\VistA\Imaging\ViX Installer for future reference or troubleshooting.) This will start the installation process. It will also start the Tomcat and Viewer/Render services.



VIX Service Installation Wizard 30.185.8.6060

1. The VIX will start automatically when installation is complete.
2. Click **Finish**. The VIX Installation Wizard will close.

VIX Service Installation Wizard 30.177.5.6044



VIX Service Installation Wizard 30.185.8.6060


1. See [*Verifying VIX Operations*](#_bookmark24)


# Post-installation

## McAfee Exclusions

Exclude the following directories using McAfee On Access Scan Properties:

* 1. C:\Program Files\Apache Software Foundation\Tomcat 8.0\logs
  2. C:\Program Files\VistA\Imaging\VIX.Render.Service\log
  3. C:\Program Files\VistA\Imaging\VIX.Viewer.Service\log
  4. C:\VixConfig\logs
  5. <cache drive>:\VixCache
  6. <cache drive>:\VIXRenderCache





## Verifying VIX Operations

After a new VIX is installed and is registered with the Image Exchange Service, Clinical Display (MAG\*3.0\*93 or later) and VistARad (MAG\*3.0\*90 or later) workstations will automatically start using the VIX.

* Clinical Display will begin sending all of its requests for remote data to the VIX immediately. No configuration changes are required for Clinical Display to start using the VIX.
* VistARad will need some local configuration changes to enable some of its VIX- supported capabilities; refer to the VistARad documentation for details.

**Verifying Access to the VIX Transaction Log**

VIX administrators can use the VIX transaction log to monitor VIX activities. If the transaction log can be accessed, the VIX is running.

To access the transaction log you will need:

* A VistA account that has the MAG VIX ADMIN security key assigned to it (while the log is a Web page, the VIX uses a VistA account to secure the log).
* Access to http://<FQDN>:8080/Vix/secure/VixLog.jsp

(Where <FQDN> is either the fully qualified domain name of the server the VIX is installed on.)

If you cannot access the transaction log, verify that the VIX service is running.

If the VIX is running but you cannot access the transaction log, ensure that port 8080 on the VIX server is not blocked. Possible culprits of a blocked port include antivirus firewalls and modifications to ACLs (Access Control Lists).


For detailed information about the contents of the transaction log, refer to the [*VIX*](https://www.va.gov/vdl/application.asp?appid=105)

[*Administrator’s Guide.*](https://www.va.gov/vdl/application.asp?appid=105)

You can also spot-check individual remote images to see if they were delivered via the VIX as described in this section.


**Spot-checking VIX Image Delivery**

1. On the Clinical Display workstation, select a patient with remote images.
2. If it is not visible already, display the Abstracts area to display an abstract for one of the remote images.
3. Right-click the abstract for the remote image and open the Image Information window.
4. In the Image Information window, check the IEN (Internal Entry Number) value. If the value starts with “urn”, the remote image was retrieved by the VIX.
5. Go to the VIX homepage: http://<FQDN of VIX server>:8080/
6. The current version is listed in the format XX.XXX.X.X. The first two digits represent Version 3.0 of the VistA Imaging system and do not change. The next three digits are the number of the latest patch that has affected the VIX. This patch number should match the number of the VIX component you have most recently installed.




Welcome to the VISA Service, your current version is 30.185.8.6060

1. Verify VIX Viewer and VIX Render services are running.



1. Verify listener tool is running.



## Using the VIX Installation Wizard to Reconfigure the VIX

You will need to re-execute the VIX Installation wizard if you need to:

* + Change the drive where the VixCache or VixConfig folders are located. It is recommended that these folders reside on the same shared drive.
  + Change the VIX configuration to use a different local VistA host name or port number.

Using the VIX Installation wizard is broadly similar to the steps for updating a VIX, except the process is much faster since no actual software is being installed.

**Note:** The following steps that new cache location is set up and/or that the local VistA database connection change has already been made.


**Tip:** Changing a VIX cache location or local connection information should take about 5 minutes.

**Reconfiguring a VIX – Standalone Server**

1. Review the information in the [*Scheduling Downtime and Impact of a VIX Update*](#_bookmark19)

section, and schedule downtime and notify appropriate groups of the downtime.

1. Log in as an administrator on the server where the VIX is installed and choose **Start | All Programs | VistA Imaging Programs | VIX Installation Service Wizard**.



1. Click **Next** until the Specify the VA Site... page displays.
2. In this page, verify that the **Site Number** box shows your STATION NUMBER (field (#99) in the INSTITUTION file (#4)). Then, click **Lookup Server Address**.
3. Verify that the correct hostname and port number for the local VistA system is displayed. Then, click **Next**.
4. When the Install the VIX Prerequisites page displays, click **Next**. (All prerequisites are installed already).
5. In the Specify the location... page, do one of the following:
   * If you are changing the location of the VIX cache and configuration files, select the new drive for each (the same drive should be used). Then, click **Create**. Then, click **Next**.
   * If you are NOT changing the location of the VIX cache and configuration files, click **Next**.
6. In the Specify the Release of Information (ROI) Configuration, do the following:
   * If you want to change any of the values, enter the new values. If you are changing the email address or addresses for invalid ROI periodic processing credentials notification, click **Validate**. If you do not want to change any of the values on this page, skip this step.
   * Click **Next**.
7. In the Configure local DoD connection page, click **Next**.
8. In the Install the VIX page, click **Install**.
9. Wait until the installation is complete and click **Finish**. The VIX will restart automatically.

# Troubleshooting

## Resuming an Interrupted VIX Installation


If you have had to interrupt or cancel an in-progress VIX installation, you can resume the installation.

**Note:** If you re-run MAG3\_0P<number>\_VIX\_Setup.msi, you are repeating the installation of the VIX Installation Wizard software, not the installation of the VIX itself. If you do this, click Cancel, choose **Yes** when prompted for confirmation, and then exit the installer. Then restart the VIX installation.

**Resuming Installation (single server VIX)**

1. Log onto the VIX server as an administrator.

### Click Start | All Programs | VistA Imaging Programs | VIX Installation Wizard.



1. Follow the steps for a first-time standalone VIX installation or for updating a standalone VIX.

## VIX Support

If you encounter problems installing the VIX please call the National Service Desk at 1- 855-673-4357.

This page is intentionally blank.

# Back Out/Uninstall

Back Out/Uninstall

## Back Out/Uninstall Scenarios

The information in this section addresses three possible cases that involve uninstalling a VIX:

* Troubleshooting
* Relocation onto a different server
* Decommission

Each of these scenarios is outlined in the following sections.

**Uninstall/Restore as part of Troubleshooting**

If you need to remove and then immediately reinstall the VIX on the same server for troubleshooting purposes, you will need to do the following:

1. Locate the product serial number for the Laurel Bridge software that is bundled with the VIX. This number is in the license paperwork that was provided by [REDACTED](mailto:VHAVILBLicenses@va.gov) when the VIX was set up.
2. **Note:** You will need to re-enter this serial number as a part of the VIX installation process.

### Choose Start | All Programs | VistA Imaging Programs | VIX Installation Service Wizard.

1. When the Welcome page displays, verify that the screen displays “This wizard will guide you through the installation of the Vista Imaging Patch 185 VIX and click Next.
2. Then, when you are prompted to do so, click **Uninstall version 30.185.xxx**. (The wizard will gracefully stop the VIX service before performing the uninstall.)
3. Go to the Control Panel, choose Add/Remove Programs, and remove the MAG\*3.0\*185 VIX installer.
4. Re-execute the VIX installation as if for a new standaloneVIX, see [*New VIX Installation - Standalone Server*](#_bookmark12) section. .

**Relocating a VIX onto a New Server**

If you need to remove all traces of the VIX from the old server and set up the VIX on a new server, you will need to do the following:

1. Contact the [REDACTED](mailto:VHAVILBLicenses@va.gov) mail group and arrange to have the existing Laurel Bridge DCF toolkit licenses transferred to a new server.
2. Validate the new server where the VIX will be installed as described in the [*Selecting and Validating the VIX Server*](#_bookmark8) section.
3. Manually remove the VIX as described in the [*Uninstalling the VIX*](#_bookmark38) section below.
4. Re-execute the VIX installation as if for a new standalone VIX, see [*New VIX Installation – Standalone Server*](#_bookmark12) section. .

**Decommissioning a VIX**

If a VIX to be completely decommissioned and not replaced by another VIX, do the following:

1. Notify the [REDACTED](mailto:VHAVILBLicenses@va.gov) mail group that the Laurel Bridge license seats used by your site are no longer being used.
2. Contact [REDACTED](mailto:VHAVIVIXSETUP@VA.GOV) mail group to have the VIX security certificates retired and the VIX removed from the site service.
3. Manually remove the VIX as described in the [*Uninstalling the VIX*](#_bookmark38) section below.
4. In VistA, remove the MAG VIX ADMIN security key from the accounts that have this key assigned.

## Uninstalling the VIX

The following steps explain how to completely remove a VIX and all its supporting components (toolkits, runtime environments, etc.) from the server where the VIX is installed.

These steps will completely remove the VIX and permanently delete the VIX cache.

Depending on the VIX server configuration and operating system, the specifics of removing the VIX will vary, but the general process is as follows:

1. Stop the VIX service.
2. Remove VIX-related applications, accounts, directories, and variables

These steps do not require a server reboot, and can be performed while the rest of the Imaging system is active.

**Stopping the VIX service**

The steps for stopping the VIX service vary based on operating system.

**Remove VIX-related applications, accounts, directories, and variables**

After stopping (and/or removing) the VIX service as described in the previous sections, you will need to remove VIX software and settings.

These steps presume you are already logged in as an administrator on the server where the VIX service used to reside.

These steps cover all supported VIX configurations.

1. Use Windows Explorer to navigate to the drive where the VIX Cache is located, and delete the following folders:

<shared drive letter>\VixCache

<shared drive letter>\VixConfig

1. Use Windows Explorer to delete the following directories: C:\ DCF\_RunTime

C:\Program Files\Java\jre1.8.0\_121

1. If you are removing the VIX permanently, also delete the following folders:

SKIP THIS STEP if you are uninstalling and reinstalling the VIX on the same server for troubleshooting purposes. If you delete these folders, you will need to recreate the VIX configuration and request new VIX security certificates.

<shared drive letter>\VixConfig C:\VixCertStore

1. Open the window used to remove programs.
   * On Windows click Start | Control Panel. Under the Programs item, click

### Uninstall a program.

* + Remove these three programs (no reboot required): Apache Tomcat 8.0.45

Java (TM) 1.8.0. Update 121

VIX Service Installation Wizard

1. In Windows Explorer, right-click the Local Disk (C:) folder, select **Properties**. Then, select the Security tab.
2. Select the apachetomcat user and click **Remove**.
3. Click **OK** to close the Properties dialog box, and click **Yes** when are asked if you want to continue.

**Note:** If one or more “Error applying security” messages display, click **Continue**

until they are all closed.

1. Open the Computer Management/Server Manager window.
   * On Windows, right-click Computer on the desktop. Then, click **Manage**.
2. In the tree on the left side of the window, navigate to Users.
   * On Windows, go to Server Manager/Configuration/ Local Users and Groups/Users.
3. In the right side of the window, right-click the apachetomcat user, click **Delete**, and click **Yes** when you are asked for confirmation.
4. Open the System Properties dialog box.
   * On Windows, right-click Computer on the desktop. Then, click **Properties**. Then on the left side of the System window, click **Advanced system settings**.
5. In the Advanced tab, click **Environment Variables**.
6. In the System variables list near the bottom of the dialog, delete the following variables:

|  |  |
| --- | --- |
| CATALINA\_HOME DCF\_BIN DCF\_CFG DCF\_CLASSES DCF\_LIB DCF\_LOG DCF\_PLATFORM DCF\_ROOT DCF\_TMP  DCF\_USER\_BIN | DCF\_USER\_CLASSES DCF\_USER\_LIB DCF\_USER\_ROOT LD\_LIBRARY\_PATH OMNI\_BIN OMNI\_LIB OMNI\_ROOT  vixcache vixconfig |

1. In the System variables list, select the Path system variable. Then, click **Edit**.
2. In the Variable value box, delete the following substrings: C:\DCF (if present)

C:\DCF\_Runtime\bin C:\DCF\_Runtime\lib

C:\Program Files\Java\jre1.8.0\_xx\bin

**Note** It is recommend that after deleting each substring you delete any extra semicolon characters.

1. After removing the substrings, click **OK**. Then click **OK** twice more to close the Environment Variables and System Properties dialog boxes.
2. If the VIX is installed on a standalone server, VIX removal is complete.

This page intentionally left blank.

# Appendix A: VIX Install Checklist

Appendix A: VIX Install Checklist

The checklist on this page summarizes the VIX installation process.

## VIX Install Checklist

|  |  |  |
| --- | --- | --- |
|  | **Requirement** | **Action** |
| ☐ | Install patch 185 KIDS in VistA | See Patch Description Document |
| ☐ | Verify VIX server/VM meets minimum hardware specifications | Allocate more CPU cores/RAM as needed to meet minimum or recommended specifications |
| ☐ | Verify Microsoft Visual C++ 2010 Redistributable Package (x64) is installed | Install Microsoft Visual C++ 2010 Redistributable Package (x64) if required. (Download from SFTP site) |
| ☐ | Verify .NET 3.5 is installed and enabled | Install and enable .NET 3.5 if required |
| ☐ | Verify .NET 4.5 is installed and enabled | Install and enable .NET 4.5 if required (Download from SFTP site) |
| ☐ | Verify you have the SQLExpress\_x64-12\_0\_2000\_8.zip file | Download from SFTP site if required |
| ☐ | Initiate the installation of the 185 VIX | See Installation Guide |
| ☐ | Verify if install is complete | See Installation Guide |
| ☐ | Verify that VIX is running | See Installation Guide |
| ☐ | Verify that VIX Viewer and VIX render services started | Verify in window services console |
| ☐ | Verify ListenerTool service is running | Verify in window services console |
| ☐ | Verify that MacAfee Exclusions are Updated | See Installation Guide |
| ☐ | Request Site Service update to your facility if needed | Enter in national CA ticket requesting CLIN3 to update your site service entry to  include the new viewer. (Port 343) |
| ☐ | Verify that images can be viewed via zero-footprint Image Viewer. | Verify JLV or other application utilizing  zero footprint viewer. If unable to view an image, enter a support ticket for assistance. |